## Sample Report
row | id1 | rt | ampl | area | comp_class | size_group | dD_raw | dD_calibrated
3 | CA02 | 1645.2 | 2379 | 9.899 | C25 Alkane | Normal | -195.462 | -203.7
3 | CA02 | 1799.1 | 6967 | 37.45 | C27 Alkane | Normal | -218.085 | -203
3 | CA02 | 1938.7 | 4441 | 22.398 | C29 Alkane | Normal | -203.052 | -197.4
3 | CA02 | 2069.9 | 2392 | 11.626 | C31 Alkane | Normal | -180.902 | -187.5
3 | CA02 | 2211.6 | 1079 | 5.461 | C33 Alkane | Normal | -161 | -183.3
4 | CA02 | 1646.7 | 2187 | 9.043 | C25 Alkane | Normal | -194.562 | -204.8
4 | CA02 | 1800.3 | 6449 | 34.264 | C27 Alkane | Normal | -216.446 | -203
4 | CA02 | 1939.9 | 4095 | 20.321 | C29 Alkane | Normal | -202.346 | -198.5
4 | CA02 | 2071.4 | 2169 | 10.436 | C31 Alkane | Normal | -178.819 | -187.4
4 | CA02 | 2213.3 | 959 | 4.845 | C33 Alkane | Normal | -157.104 | -181.6
5 | CA10 | 1646.5 | 2214 | 9.508 | C25 Alkane | Normal | -192.959 | -202.7
5 | CA10 | 1799.7 | 5567 | 29.008 | C27 Alkane | Normal | -212.532 | -202.3
5 | CA10 | 1939.9 | 4262 | 21.397 | C29 Alkane | Normal | -200.238 | -195.3
5 | CA10 | 2072.2 | 3384 | 17.215 | C31 Alkane | Normal | -187.56 | -186.9
5 | CA10 | 2214.1 | 1864 | 9.897 | C33 Alkane | Normal | -177.288 | -189.2
6 | CA10 | 1646.3 | 1892 | 7.742 | C25 Alkane | Normal | -189.437 | -202.5
6 | CA10 | 1799.1 | 4775 | 23.3 | C27 Alkane | Normal | -209.964 | -203.1
6 | CA10 | 1939.3 | 3433 | 16.693 | C29 Alkane | Normal | -196.246 | -196
6 | CA10 | 2071.4 | 2644 | 12.987 | C31 Alkane | Normal | -184.328 | -188.9
6 | CA10 | 2213.3 | 1395 | 7.106 | C33 Alkane | Normal | -172.499 | -190.5
7 | CA14 | 1646.3 | 2258 | 9.417 | C25 Alkane | Normal | -194.191 | -203.6
7 | CA14 | 1800.3 | 6948 | 38.063 | C27 Alkane | Normal | -218.097 | -203
7 | CA14 | 1939.9 | 4800 | 24.634 | C29 Alkane | Normal | -204.403 | -197.1
7 | CA14 | 2071.6 | 2995 | 15.031 | C31 Alkane | Normal | -182.882 | -184.5
7 | CA14 | 2213.3 | 1478 | 7.566 | C33 Alkane | Normal | -168.132 | -184.2
8 | CA14 | 1646.7 | 2258 | 9.41 | C25 Alkane | Normal | -194.338 | -203.7
8 | CA14 | 1800.7 | 6972 | 38.03 | C27 Alkane | Normal | -217.2 | -202
8 | CA14 | 1940.4 | 4814 | 24.577 | C29 Alkane | Normal | -203.806 | -196.4
8 | CA14 | 2071.8 | 2997 | 15.037 | C31 Alkane | Normal | -180.947 | -182.4
8 | CA14 | 2213.7 | 1471 | 7.589 | C33 Alkane | Normal | -168.529 | -184.7
9 | CA18 | 1647.3 | 3311 | 14.501 | C25 Alkane | Normal | -198.246 | -199
9 | CA18 | 1801.2 | 8290 | 47.778 | C27 Alkane | Large | -218.374 | -199.3
9 | CA18 | 1940.4 | 5313 | 27.739 | C29 Alkane | Normal | -203.489 | -193.8
9 | CA18 | 2071.8 | 3121 | 15.752 | C31 Alkane | Normal | -179.613 | -180
9 | CA18 | 2213.7 | 1728 | 8.989 | C33 Alkane | Normal | -168.166 | -180.6
10 | CA18 | 1647.3 | 3325 | 14.524 | C25 Alkane | Normal | -197.325 | -197.9
10 | CA18 | 1801.4 | 8294 | 47.713 | C27 Alkane | Large | -218.138 | -199.1
10 | CA18 | 1940.6 | 5302 | 27.7 | C29 Alkane | Normal | -203.422 | -193.8
10 | CA18 | 2071.8 | 3128 | 15.701 | C31 Alkane | Normal | -179.634 | -180
10 | CA18 | 2213.9 | 1708 | 8.967 | C33 Alkane | Normal | -170.322 | -183.3
27 | BOF1-11 | 1483.5 | 1362 | 5.271 | C23 Alkane | Normal | -210.192 | -234.4
27 | BOF1-11 | 1649.4 | 5955 | 29.805 | C25 Alkane | Normal | -251.63 | -241.6
27 | BOF1-11 | 1802.6 | 9300 | 56.28 | C27 Alkane | Large | -257.327 | -236.3
27 | BOF1-11 | 1940.4 | 3981 | 19.684 | C29 Alkane | Normal | -227.744 | -226.4
27 | BOF1-11 | 2072.9 | 3541 | 17.978 | C31 Alkane | Normal | -224.135 | -225.5
28 | BOF1-11 | 1483.1 | 1375 | 5.323 | C23 Alkane | Normal | -212.613 | -237
28 | BOF1-11 | 1649.2 | 5954 | 30.181 | C25 Alkane | Normal | -252.168 | -242.2
28 | BOF1-11 | 1802.2 | 9416 | 57.292 | C27 Alkane | Large | -257.123 | -235.8
28 | BOF1-11 | 1940.1 | 4064 | 20.34 | C29 Alkane | Normal | -227.801 | -226
28 | BOF1-11 | 2072.4 | 3711 | 18.929 | C31 Alkane | Normal | -226.587 | -227
29 | BOF1-31 | 1484.1 | 3143 | 13.154 | C23 Alkane | Normal | -239.168 | -244.9
29 | BOF1-31 | 1648.6 | 5726 | 28.615 | C25 Alkane | Normal | -248.482 | -239.3
29 | BOF1-31 | 1800.1 | 6452 | 34.326 | C27 Alkane | Normal | -245.279 | -233
29 | BOF1-31 | 1939.7 | 4254 | 21.191 | C29 Alkane | Normal | -235.514 | -233.1
29 | BOF1-31 | 2071.6 | 3204 | 16.026 | C31 Alkane | Normal | -231.464 | -236
29 | BOF1-31 | 2213.1 | 1051 | 5.246 | C33 Alkane | Normal | -208.194 | -239
30 | BOF1-31 | 1484.7 | 3112 | 13.015 | C23 Alkane | Normal | -240.459 | -246.5
30 | BOF1-31 | 1649.2 | 5707 | 28.198 | C25 Alkane | Normal | -249.167 | -240.1
30 | BOF1-31 | 1800.7 | 6293 | 33.553 | C27 Alkane | Normal | -244.905 | -233.2
30 | BOF1-31 | 1940.4 | 4136 | 20.41 | C29 Alkane | Normal | -235.896 | -234.2
30 | BOF1-31 | 2072.2 | 3024 | 15.101 | C31 Alkane | Normal | -230.009 | -235.9
30 | BOF1-31 | 2213.5 | 976 | 4.822 | C33 Alkane | Normal | -206.58 | -239.1
47 | BOF1-41 | 1644.6 | 1930 | 8.853 | C25 Alkane | Normal | -216.186 | -232.1
47 | BOF1-41 | 1795.7 | 1985 | 9.55 | C27 Alkane | Normal | -212.574 | -227.3
47 | BOF1-41 | 1936.4 | 1822 | 9.016 | C29 Alkane | Normal | -219.17 | -237
48 | BOF1-41 | 1482.4 | 1363 | 5.278 | C23 Alkane | Normal | -211.056 | -235.4
48 | BOF1-41 | 1646.1 | 2474 | 10.462 | C25 Alkane | Normal | -223.113 | -233.4
48 | BOF1-41 | 1797.4 | 2554 | 11.376 | C27 Alkane | Normal | -218.251 | -227.2
48 | BOF1-41 | 1938.1 | 2395 | 11.045 | C29 Alkane | Normal | -225.166 | -236.5
49 | NE 0-5 cm | 1796.6 | 1995 | 8.538 | C27 Alkane | Normal | -223.958 | -240
49 | NE 0-5 cm | 1939.3 | 4910 | 25.073 | C29 Alkane | Normal | -242.186 | -236.6
49 | NE 0-5 cm | 2072.2 | 5406 | 30.143 | C31 Alkane | Normal | -250.052 | -242.5
49 | NE 0-5 cm | 2215.2 | 3876 | 23.059 | C33 Alkane | Normal | -239.154 | -239.4
50 | NE 0-5 cm | 1797.4 | 2004 | 8.498 | C27 Alkane | Normal | -226.074 | -242.2
50 | NE 0-5 cm | 1940.1 | 4899 | 25.075 | C29 Alkane | Normal | -243.314 | -237.9
50 | NE 0-5 cm | 2073.1 | 5422 | 30.148 | C31 Alkane | Normal | -250.906 | -243.3
50 | NE 0-5 cm | 2216.2 | 3898 | 23.089 | C33 Alkane | Normal | -239.276 | -239.4

## Standard Report
std_mix | conc | comp_class | ungrouped_full_rmse | conc_grouped_rmse | std_mix_rmse | observed_value | observed_sd | observed_n | accepted_value | root_mean_square_error | mean_signed_difference
ACAL | 5 | C25 Alkane | 3.8 | 4.3 | 6.7 | -256.2 | 1.2 | 2 | -263.6 | 10.6 | -14.9
ACAL | 5 | C27 Alkane | 3.8 | 4.3 | 6.7 | -174.3 | 2.7 | 2 | -172.8 | 3.4 | 2.9
ACAL | 5 | C29 Alkane | 3.8 | 4.3 | 6.7 | -92.1 | 3 | 2 | -85.4 | 10 | 13.5
ACAL | 10 | C25 Alkane | 3.8 | 3.3 | 4.8 | -257.6 | 0.5 | 2 | -263.6 | 8.5 | -12
ACAL | 10 | C27 Alkane | 3.8 | 3.3 | 4.8 | -176 | 0.4 | 2 | -172.8 | 4.5 | 6.4
ACAL | 10 | C29 Alkane | 3.8 | 3.3 | 4.8 | -88.7 | 1 | 2 | -85.4 | 4.7 | 6.6
ACAL | 20 | C25 Alkane | 3.8 | 3.8 | 4.6 | -257.5 | 3.6 | 12 | -263.6 | 7.3 | -6.7
ACAL | 20 | C27 Alkane | 3.8 | 3.8 | 4.6 | -174.8 | 2.3 | 12 | -172.8 | 3.1 | 2.2
ACAL | 20 | C29 Alkane | 3.8 | 3.8 | 4.6 | -87.3 | 1.2 | 12 | -85.4 | 2.3 | 2.1
C4 | 5 | C17 Alkane | 3.8 | 4.3 | 2.5 | -119.9 | 1.4 | 2 | -117.8 | 3.2 | 4.1
C4 | 5 | C19 Alkane | 3.8 | 4.3 | 2.5 | -55.1 | 1 | 2 | -56.3 | 2 | -2.4
C4 | 5 | C21 Alkane | 3.8 | 4.3 | 2.5 | -181.5 | 0.7 | 2 | -177.8 | 5.3 | 7.5
C4 | 5 | C23 Alkane | 3.8 | 4.3 | 2.5 | -68.6 | 2.8 | 2 | -67.2 | 3.4 | 2.7
C4 | 5 | C25 Alkane | 3.8 | 4.3 | 2.5 | -254.1 | 0.6 | 2 | -254.1 | 0.6 | -0.1
C4 | 10 | C17 Alkane | 3.8 | 3.3 | 2.2 | -118.5 | 1.1 | 2 | -117.8 | 1.5 | 1.4
C4 | 10 | C19 Alkane | 3.8 | 3.3 | 2.2 | -52.3 | 0.9 | 2 | -56.3 | 5.7 | -7.9
C4 | 10 | C21 Alkane | 3.8 | 3.3 | 2.2 | -179.8 | 0.1 | 2 | -177.8 | 2.9 | 4.1
C4 | 10 | C23 Alkane | 3.8 | 3.3 | 2.2 | -67.3 | 0.2 | 2 | -67.2 | 0.2 | 0.2
C4 | 10 | C25 Alkane | 3.8 | 3.3 | 2.2 | -254.9 | 0.6 | 2 | -254.1 | 1.2 | 1.5
C4 | 20 | C17 Alkane | 3.8 | 3.8 | 1.5 | -119.1 | 0.8 | 4 | -117.8 | 1.7 | 1.7
C4 | 20 | C19 Alkane | 3.8 | 3.8 | 1.5 | -53.7 | 0.6 | 4 | -56.3 | 3.1 | -3.5
C4 | 20 | C21 Alkane | 3.8 | 3.8 | 1.5 | -178.6 | 0.1 | 4 | -177.8 | 1 | 1.1
C4 | 20 | C23 Alkane | 3.8 | 3.8 | 1.5 | -66.9 | 0.6 | 4 | -67.2 | 0.7 | -0.4
C4 | 20 | C25 Alkane | 3.8 | 3.8 | 1.5 | -253.8 | 0.5 | 4 | -254.1 | 0.6 | -0.4

## Sample Info
id1 | sample_type | unique_id | target_peak | target_peak_conc | volume | ng_target_peak
CA02 |  |  |  |  | 1350 | 27000
CA10 |  |  |  |  | 347.7 | 6954
CA14 |  |  |  |  | 202.3 | 4046
CA18 |  |  |  |  | 86.35 | 1727
ACAL
ACAL
BOF1-11 |  |  |  |  | 373.15 | 7463
BOF1-31 |  |  |  |  | 399.3 | 7986
C4
C4
BOF1-41 |  |  |  |  | 18.65 | 373
NE 0-5 cm |  |  |  |  | 397.7163 | 7954.325

## Final Processing Table
row | id1 | id2 | conc | rt | raw_R | dD_raw | row_base | comp_class | comp | class | std_mix | relative_concentration | dD_known | dD_known_sd | dD_known_n | dD_processing | method | drift_correction | dD_predrift | dD_drift | acceptable_peak_value | size_value | size_lower | size_upper | size_group | ampl | area | dD_presize | dD_stage1_correction | size_slope | size_intercept | dD_size | i | normalization_method | intercept | slope | dD_prenormalization | dD_normalization | bound_hydrogen_count | derivative_hydrogen_count | derivative_dD | derivative_dD_uncertainty | total_hydrogen_count | final_dD
1 | ACAL | standard_control_drift | 20 | 1650.5 | 0.26209 | -261.628 | 1 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -263.9784 | 3 | 0 | -261.628 | -261.628 | 3820 | 3820 | 933 | 7189 | Normal | 3820 | 17.064 | -261.628 | -270.8865 | -25.4376 | 220.1491 | -271.9676 | 1 | 1 | 25.199 | 1.0633 | -271.9676 | -263.9784 | 1 | 0 | -192 | 0 | 1 | -263.9784
1 | ACAL | standard_control_drift | 20 | 1800.5 | 0.28992 | -183.241 | 1 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -179.793 | 3 | 0 | -183.241 | -183.241 | 3775 | 3775 | 933 | 7189 | Normal | 3775 | 17.63 | -183.241 | -192.7217 | -20.3176 | 176.8906 | -192.7923 | 1 | 1 | 25.199 | 1.0633 | -192.7923 | -179.793 | 1 | 0 | -192 | 0 | 1 | -179.793
1 | ACAL | standard_control_drift | 20 | 1941 | 0.31965 | -99.491 | 1 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -89.0914 | 3 | 0 | -99.491 | -99.491 | 3873 | 3873 | 933 | 7189 | Normal | 3873 | 19.056 | -99.491 | -108.4911 | -14.8004 | 130.2752 | -107.4886 | 1 | 1 | 25.199 | 1.0633 | -107.4886 | -89.0914 | 1 | 0 | -192 | 0 | 1 | -89.0914
2 | ACAL | standard_control_drift | 20 | 1649.8 | 0.26187 | -262.098 | 1 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -264.6427 | 3 | 0 | -262.098 | -262.098 | 3798 | 3798 | 933 | 7189 | Normal | 3798 | 17.028 | -262.098 | -271.4648 | -25.4755 | 220.4692 | -272.5924 | 1 | 1 | 25.199 | 1.0633 | -272.5924 | -264.6427 | 1 | 0 | -192 | 0 | 1 | -264.6427
2 | ACAL | standard_control_drift | 20 | 1799.9 | 0.29001 | -182.804 | 1 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -179.2122 | 3 | 0 | -182.804 | -182.804 | 3794 | 3794 | 933 | 7189 | Normal | 3794 | 17.614 | -182.804 | -192.1906 | -20.2829 | 176.5967 | -192.2461 | 1 | 1 | 25.199 | 1.0633 | -192.2461 | -179.2122 | 1 | 0 | -192 | 0 | 1 | -179.2122
2 | ACAL | standard_control_drift | 20 | 1940.4 | 0.31929 | -100.29 | 1 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -89.8565 | 3 | 0 | -100.29 | -100.29 | 3896 | 3896 | 933 | 7189 | Normal | 3896 | 19.038 | -100.29 | -109.1791 | -14.8454 | 130.6559 | -108.2082 | 1 | 1 | 25.199 | 1.0633 | -108.2082 | -89.8565 | 1 | 0 | -192 | 0 | 1 | -89.8565
3 | CA02 | sample |  | 1645.2 | 0.28532 | -195.462 | 3 | C25 Alkane | C25 | Alkane | [5 empty] | -203.7429 | 3 | 0 | -195.462 | -195.462 | 2379 | 2379 | 933 | 7189 | Normal | 2379 | 9.899 | -195.462 | -213.6002 | -21.6852 | 188.4454 | -215.3169 | 1 | 1 | 25.199 | 1.0633 | -215.3169 | -203.7429 | 1 | 0 | -192 | 0 | 1 | -203.7429
3 | CA02 | sample |  | 1799.1 | 0.2773 | -218.085 | 3 | C27 Alkane | C27 | Alkane | [5 empty] | -202.9531 | 3 | 0 | -218.085 | -218.085 | 6967 | 6967 | 933 | 7189 | Normal | 6967 | 37.45 | -218.085 | -216.0756 | -21.8474 | 189.8153 | -214.5742 | 1 | 1 | 25.199 | 1.0633 | -214.5742 | -202.9531 | 1 | 0 | -192 | 0 | 1 | -202.9531
3 | CA02 | sample |  | 1938.7 | 0.28263 | -203.052 | 3 | C29 Alkane | C29 | Alkane | [5 empty] | -197.4063 | 3 | 0 | -203.052 | -203.052 | 4441 | 4441 | 933 | 7189 | Normal | 4441 | 22.398 | -203.052 | -209.4861 | -21.4158 | 186.1685 | -209.3574 | 1 | 1 | 25.199 | 1.0633 | -209.3574 | -197.4063 | 1 | 0 | -192 | 0 | 1 | -197.4063
3 | CA02 | sample |  | 2069.9 | 0.29049 | -180.902 | 3 | C31 Alkane | C31 | Alkane | [5 empty] | -187.4526 | 3 | 0 | -180.902 | -180.902 | 2392 | 2392 | 933 | 7189 | Normal | 2392 | 11.626 | -180.902 | -198.938 | -20.7248 | 180.3309 | -199.9961 | 1 | 1 | 25.199 | 1.0633 | -199.9961 | -187.4526 | 1 | 0 | -192 | 0 | 1 | -187.4526
3 | CA02 | sample |  | 2211.6 | 0.29754 | -161 | 3 | C33 Alkane | C33 | Alkane | [5 empty] | -183.3265 | 3 | 0 | -161 | -161 | 1079 | 1079 | 933 | 7189 | Normal | 1079 | 5.461 | -161 | -193.9633 | -20.399 | 177.5777 | -196.1156 | 1 | 1 | 25.199 | 1.0633 | -196.1156 | -183.3265 | 1 | 0 | -192 | 0 | 1 | -183.3265
4 | CA02 | sample |  | 1646.7 | 0.28584 | -194.562 | 3 | C25 Alkane | C25 | Alkane | [5 empty] | -204.762 | 3 | 0 | -194.562 | -194.562 | 2187 | 2187 | 933 | 7189 | Normal | 2187 | 9.043 | -194.562 | -214.2781 | -21.7296 | 188.8205 | -216.2754 | 1 | 1 | 25.199 | 1.0633 | -216.2754 | -204.762 | 1 | 0 | -192 | 0 | 1 | -204.762
4 | CA02 | sample |  | 1800.3 | 0.27807 | -216.446 | 3 | C27 Alkane | C27 | Alkane | [5 empty] | -203.0094 | 3 | 0 | -216.446 | -216.446 | 6449 | 6449 | 933 | 7189 | Normal | 6449 | 34.264 | -216.446 | -215.8853 | -21.8349 | 189.71 | -214.6271 | 1 | 1 | 25.199 | 1.0633 | -214.6271 | -203.0094 | 1 | 0 | -192 | 0 | 1 | -203.0094
4 | CA02 | sample |  | 1939.9 | 0.28307 | -202.346 | 3 | C29 Alkane | C29 | Alkane | [5 empty] | -198.5101 | 3 | 0 | -202.346 | -202.346 | 4095 | 4095 | 933 | 7189 | Normal | 4095 | 20.321 | -202.346 | -210.301 | -21.4691 | 186.6195 | -210.3955 | 1 | 1 | 25.199 | 1.0633 | -210.3955 | -198.5101 | 1 | 0 | -192 | 0 | 1 | -198.5101
4 | CA02 | sample |  | 2071.4 | 0.29142 | -178.819 | 3 | C31 Alkane | C31 | Alkane | [5 empty] | -187.3811 | 3 | 0 | -178.819 | -178.819 | 2169 | 2169 | 933 | 7189 | Normal | 2169 | 10.436 | -178.819 | -198.69 | -20.7086 | 180.1937 | -199.9288 | 1 | 1 | 25.199 | 1.0633 | -199.9288 | -187.3811 | 1 | 0 | -192 | 0 | 1 | -187.3811
4 | CA02 | sample |  | 2213.3 | 0.29913 | -157.104 | 3 | C33 Alkane | C33 | Alkane | [5 empty] | -181.5554 | 3 | 0 | -157.104 | -157.104 | 959 | 959 | 933 | 7189 | Normal | 959 | 4.845 | -157.104 | -192.278 | -20.2886 | 176.645 | -194.4498 | 1 | 1 | 25.199 | 1.0633 | -194.4498 | -181.5554 | 1 | 0 | -192 | 0 | 1 | -181.5554
5 | CA10 | sample |  | 1646.5 | 0.28646 | -192.959 | 5 | C25 Alkane | C25 | Alkane | [5 empty] | -202.6788 | 3 | 0 | -192.959 | -192.959 | 2214 | 2214 | 933 | 7189 | Normal | 2214 | 9.508 | -192.959 | -212.445 | -21.6096 | 187.806 | -214.3161 | 1 | 1 | 25.199 | 1.0633 | -214.3161 | -202.6788 | 1 | 0 | -192 | 0 | 1 | -202.6788
5 | CA10 | sample |  | 1799.7 | 0.27951 | -212.532 | 5 | C27 Alkane | C27 | Alkane | [5 empty] | -202.2763 | 3 | 0 | -212.532 | -212.532 | 5567 | 5567 | 933 | 7189 | Normal | 5567 | 29.008 | -212.532 | -214.7289 | -21.7592 | 189.07 | -213.9376 | 1 | 1 | 25.199 | 1.0633 | -213.9376 | -202.2763 | 1 | 0 | -192 | 0 | 1 | -202.2763
5 | CA10 | sample |  | 1939.9 | 0.28387 | -200.238 | 5 | C29 Alkane | C29 | Alkane | [5 empty] | -195.338 | 3 | 0 | -200.238 | -200.238 | 4262 | 4262 | 933 | 7189 | Normal | 4262 | 21.397 | -200.238 | -207.4435 | -21.282 | 185.0381 | -207.4122 | 1 | 1 | 25.199 | 1.0633 | -207.4122 | -195.338 | 1 | 0 | -192 | 0 | 1 | -195.338
5 | CA10 | sample |  | 2072.2 | 0.28837 | -187.56 | 5 | C31 Alkane | C31 | Alkane | [5 empty] | -186.8903 | 3 | 0 | -187.56 | -187.56 | 3384 | 3384 | 933 | 7189 | Normal | 3384 | 17.215 | -187.56 | -199.0909 | -20.7348 | 180.4155 | -199.4673 | 1 | 1 | 25.199 | 1.0633 | -199.4673 | -186.8903 | 1 | 0 | -192 | 0 | 1 | -186.8903
5 | CA10 | sample |  | 2214.1 | 0.29202 | -177.288 | 5 | C33 Alkane | C33 | Alkane | [5 empty] | -189.1738 | 3 | 0 | -177.288 | -177.288 | 1864 | 1864 | 933 | 7189 | Normal | 1864 | 9.897 | -177.288 | -200.0005 | -20.7944 | 180.9189 | -201.6149 | 1 | 1 | 25.199 | 1.0633 | -201.6149 | -189.1738 | 1 | 0 | -192 | 0 | 1 | -189.1738
6 | CA10 | sample |  | 1646.3 | 0.28754 | -189.437 | 5 | C25 Alkane | C25 | Alkane | [5 empty] | -202.5089 | 3 | 0 | -189.437 | -189.437 | 1892 | 1892 | 933 | 7189 | Normal | 1892 | 7.742 | -189.437 | -211.87 | -21.5719 | 187.4878 | -214.1564 | 1 | 1 | 25.199 | 1.0633 | -214.1564 | -202.5089 | 1 | 0 | -192 | 0 | 1 | -202.5089
6 | CA10 | sample |  | 1799.1 | 0.28026 | -209.964 | 5 | C27 Alkane | C27 | Alkane | [5 empty] | -203.0959 | 3 | 0 | -209.964 | -209.964 | 4775 | 4775 | 933 | 7189 | Normal | 4775 | 23.3 | -209.964 | -215.0384 | -21.7794 | 189.2413 | -214.7084 | 1 | 1 | 25.199 | 1.0633 | -214.7084 | -203.0959 | 1 | 0 | -192 | 0 | 1 | -203.0959
6 | CA10 | sample |  | 1939.3 | 0.28513 | -196.246 | 5 | C29 Alkane | C29 | Alkane | [5 empty] | -195.9894 | 3 | 0 | -196.246 | -196.246 | 3433 | 3433 | 933 | 7189 | Normal | 3433 | 16.693 | -196.246 | -207.5074 | -21.2861 | 185.0734 | -208.0249 | 1 | 1 | 25.199 | 1.0633 | -208.0249 | -195.9894 | 1 | 0 | -192 | 0 | 1 | -195.9894
6 | CA10 | sample |  | 2071.4 | 0.28936 | -184.328 | 5 | C31 Alkane | C31 | Alkane | [5 empty] | -188.9495 | 3 | 0 | -184.328 | -184.328 | 2644 | 2644 | 933 | 7189 | Normal | 2644 | 12.987 | -184.328 | -200.4859 | -20.8262 | 181.1876 | -201.4039 | 1 | 1 | 25.199 | 1.0633 | -201.4039 | -188.9495 | 1 | 0 | -192 | 0 | 1 | -188.9495
6 | CA10 | sample |  | 2213.3 | 0.29355 | -172.499 | 5 | C33 Alkane | C33 | Alkane | [5 empty] | -190.5443 | 3 | 0 | -172.499 | -172.499 | 1395 | 1395 | 933 | 7189 | Normal | 1395 | 7.106 | -172.499 | -200.646 | -20.8367 | 181.2762 | -202.9038 | 1 | 1 | 25.199 | 1.0633 | -202.9038 | -190.5443 | 1 | 0 | -192 | 0 | 1 | -190.5443
7 | CA14 | sample |  | 1646.3 | 0.28607 | -194.191 | 7 | C25 Alkane | C25 | Alkane | [5 empty] | -203.5803 | 3 | 0 | -194.191 | -194.191 | 2258 | 2258 | 933 | 7189 | Normal | 2258 | 9.417 | -194.191 | -213.308 | -21.6661 | 188.2836 | -215.164 | 1 | 1 | 25.199 | 1.0633 | -215.164 | -203.5803 | 1 | 0 | -192 | 0 | 1 | -203.5803
7 | CA14 | sample |  | 1800.3 | 0.27758 | -218.097 | 7 | C27 Alkane | C27 | Alkane | [5 empty] | -203.0276 | 3 | 0 | -218.097 | -218.097 | 6948 | 6948 | 933 | 7189 | Normal | 6948 | 38.063 | -218.097 | -216.1388 | -21.8515 | 189.8503 | -214.6442 | 1 | 1 | 25.199 | 1.0633 | -214.6442 | -203.0276 | 1 | 0 | -192 | 0 | 1 | -203.0276
7 | CA14 | sample |  | 1939.9 | 0.28244 | -204.403 | 7 | C29 Alkane | C29 | Alkane | [5 empty] | -197.0729 | 3 | 0 | -204.403 | -204.403 | 4800 | 4800 | 933 | 7189 | Normal | 4800 | 24.634 | -204.403 | -209.3795 | -21.4088 | 186.1095 | -209.0438 | 1 | 1 | 25.199 | 1.0633 | -209.0438 | -197.0729 | 1 | 0 | -192 | 0 | 1 | -197.0729
7 | CA14 | sample |  | 2071.6 | 0.29008 | -182.882 | 7 | C31 Alkane | C31 | Alkane | [5 empty] | -184.5346 | 3 | 0 | -182.882 | -182.882 | 2995 | 2995 | 933 | 7189 | Normal | 2995 | 15.031 | -182.882 | -196.7026 | -20.5784 | 179.0938 | -197.2518 | 1 | 1 | 25.199 | 1.0633 | -197.2518 | -184.5346 | 1 | 0 | -192 | 0 | 1 | -184.5346
7 | CA14 | sample |  | 2213.3 | 0.29532 | -168.132 | 7 | C33 Alkane | C33 | Alkane | [5 empty] | -184.1838 | 3 | 0 | -168.132 | -168.132 | 1478 | 1478 | 933 | 7189 | Normal | 1478 | 7.566 | -168.132 | -195.1953 | -20.4797 | 178.2596 | -196.9218 | 1 | 1 | 25.199 | 1.0633 | -196.9218 | -184.1838 | 1 | 0 | -192 | 0 | 1 | -184.1838
8 | CA14 | sample |  | 1646.7 | 0.2858 | -194.338 | 7 | C25 Alkane | C25 | Alkane | [5 empty] | -203.744 | 3 | 0 | -194.338 | -194.338 | 2258 | 2258 | 933 | 7189 | Normal | 2258 | 9.41 | -194.338 | -213.455 | -21.6757 | 188.365 | -215.318 | 1 | 1 | 25.199 | 1.0633 | -215.318 | -203.744 | 1 | 0 | -192 | 0 | 1 | -203.744
8 | CA14 | sample |  | 1800.7 | 0.27769 | -217.2 | 7 | C27 Alkane | C27 | Alkane | [5 empty] | -202.0206 | 3 | 0 | -217.2 | -217.2 | 6972 | 6972 | 933 | 7189 | Normal | 6972 | 38.03 | -217.2 | -215.1772 | -21.7885 | 189.3181 | -213.6971 | 1 | 1 | 25.199 | 1.0633 | -213.6971 | -202.0206 | 1 | 0 | -192 | 0 | 1 | -202.0206
8 | CA14 | sample |  | 1940.4 | 0.28244 | -203.806 | 7 | C29 Alkane | C29 | Alkane | [5 empty] | -196.3732 | 3 | 0 | -203.806 | -203.806 | 4814 | 4814 | 933 | 7189 | Normal | 4814 | 24.577 | -203.806 | -208.7279 | -21.3661 | 185.7489 | -208.3858 | 1 | 1 | 25.199 | 1.0633 | -208.3858 | -196.3732 | 1 | 0 | -192 | 0 | 1 | -196.3732
8 | CA14 | sample |  | 2071.8 | 0.29055 | -180.947 | 7 | C31 Alkane | C31 | Alkane | [5 empty] | -182.4024 | 3 | 0 | -180.947 | -180.947 | 2997 | 2997 | 933 | 7189 | Normal | 2997 | 15.037 | -180.947 | -194.7551 | -20.4508 | 178.016 | -195.2465 | 1 | 1 | 25.199 | 1.0633 | -195.2465 | -182.4024 | 1 | 0 | -192 | 0 | 1 | -182.4024
8 | CA14 | sample |  | 2213.7 | 0.29496 | -168.529 | 7 | C33 Alkane | C33 | Alkane | [5 empty] | -184.7484 | 3 | 0 | -168.529 | -168.529 | 1471 | 1471 | 933 | 7189 | Normal | 1471 | 7.589 | -168.529 | -195.6813 | -20.5115 | 178.5286 | -197.4528 | 1 | 1 | 25.199 | 1.0633 | -197.4528 | -184.7484 | 1 | 0 | -192 | 0 | 1 | -184.7484
9 | CA18 | sample |  | 1647.3 | 0.28445 | -198.246 | 9 | C25 Alkane | C25 | Alkane | [5 empty] | -198.9991 | 3 | 0 | -198.246 | -198.246 | 3311 | 3311 | 933 | 7189 | Normal | 3311 | 14.501 | -198.246 | -210.1858 | -21.4616 | 186.5557 | -210.8555 | 1 | 1 | 25.199 | 1.0633 | -210.8555 | -198.9991 | 1 | 0 | -192 | 0 | 1 | -198.9991
9 | CA18 | sample |  | 1801.2 | 0.27731 | -218.374 | 9 | C27 Alkane | C27 | Alkane | [5 empty] | -199.3403 | 3 | 0 | -218.374 | -218.374 | 8290 | 8290 | 933 | 7189 | Large | 8290 | 47.778 | -218.374 | -213.1046 | -21.6528 | 188.171 | -211.1763 | 1 | 1 | 25.199 | 1.0633 | -211.1763 | -199.3403 | 1 | 0 | -192 | 0 | 1 | -199.3403
9 | CA18 | sample |  | 1940.4 | 0.28259 | -203.489 | 9 | C29 Alkane | C29 | Alkane | [5 empty] | -193.8149 | 3 | 0 | -203.489 | -203.489 | 5313 | 5313 | 933 | 7189 | Normal | 5313 | 27.739 | -203.489 | -206.5616 | -21.2242 | 184.55 | -205.9798 | 1 | 1 | 25.199 | 1.0633 | -205.9798 | -193.8149 | 1 | 0 | -192 | 0 | 1 | -193.8149
9 | CA18 | sample |  | 2071.8 | 0.29106 | -179.613 | 9 | C31 Alkane | C31 | Alkane | [5 empty] | -180.0436 | 3 | 0 | -179.613 | -179.613 | 3121 | 3121 | 933 | 7189 | Normal | 3121 | 15.752 | -179.613 | -192.6609 | -20.3137 | 176.857 | -193.0281 | 1 | 1 | 25.199 | 1.0633 | -193.0281 | -180.0436 | 1 | 0 | -192 | 0 | 1 | -180.0436
9 | CA18 | sample |  | 2213.7 | 0.29512 | -168.166 | 9 | C33 Alkane | C33 | Alkane | [5 empty] | -180.6164 | 3 | 0 | -168.166 | -168.166 | 1728 | 1728 | 933 | 7189 | Normal | 1728 | 8.989 | -168.166 | -192.2991 | -20.29 | 176.6567 | -193.5667 | 1 | 1 | 25.199 | 1.0633 | -193.5667 | -180.6164 | 1 | 0 | -192 | 0 | 1 | -180.6164
10 | CA18 | sample |  | 1647.3 | 0.28487 | -197.325 | 9 | C25 Alkane | C25 | Alkane | [5 empty] | -197.8999 | 3 | 0 | -197.325 | -197.325 | 3325 | 3325 | 933 | 7189 | Normal | 3325 | 14.524 | -197.325 | -209.1857 | -21.3961 | 186.0023 | -209.8217 | 1 | 1 | 25.199 | 1.0633 | -209.8217 | -197.8999 | 1 | 0 | -192 | 0 | 1 | -197.8999
10 | CA18 | sample |  | 1801.4 | 0.27748 | -218.138 | 9 | C27 Alkane | C27 | Alkane | [5 empty] | -199.0881 | 3 | 0 | -218.138 | -218.138 | 8294 | 8294 | 933 | 7189 | Large | 8294 | 47.713 | -218.138 | -212.8595 | -21.6367 | 188.0354 | -210.9391 | 1 | 1 | 25.199 | 1.0633 | -210.9391 | -199.0881 | 1 | 0 | -192 | 0 | 1 | -199.0881
10 | CA18 | sample |  | 1940.6 | 0.2827 | -203.422 | 9 | C29 Alkane | C29 | Alkane | [5 empty] | -193.7907 | 3 | 0 | -203.422 | -203.422 | 5302 | 5302 | 933 | 7189 | Normal | 5302 | 27.7 | -203.422 | -206.5334 | -21.2223 | 184.5344 | -205.957 | 1 | 1 | 25.199 | 1.0633 | -205.957 | -193.7907 | 1 | 0 | -192 | 0 | 1 | -193.7907
10 | CA18 | sample |  | 2071.8 | 0.29114 | -179.634 | 9 | C31 Alkane | C31 | Alkane | [5 empty] | -180.017 | 3 | 0 | -179.634 | -179.634 | 3128 | 3128 | 933 | 7189 | Normal | 3128 | 15.701 | -179.634 | -192.6399 | -20.3123 | 176.8454 | -193.003 | 1 | 1 | 25.199 | 1.0633 | -193.003 | -180.017 | 1 | 0 | -192 | 0 | 1 | -180.017
10 | CA18 | sample |  | 2213.9 | 0.29445 | -170.322 | 9 | C33 Alkane | C33 | Alkane | [5 empty] | -183.3263 | 3 | 0 | -170.322 | -170.322 | 1708 | 1708 | 933 | 7189 | Normal | 1708 | 8.967 | -170.322 | -194.6734 | -20.4455 | 177.9707 | -196.1154 | 1 | 1 | 25.199 | 1.0633 | -196.1154 | -183.3263 | 1 | 0 | -192 | 0 | 1 | -183.3263
11 | ACAL | standard_control_drift | 20 | 1650.1 | 0.26359 | -257.433 | 11 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -258.6727 | 3 | 0 | -257.433 | -257.433 | 3933 | 3933 | 933 | 7189 | Normal | 3933 | 17.632 | -257.433 | -266.1449 | -25.127 | 217.525 | -266.9777 | 2 | 1 | 25.199 | 1.0633 | -266.9777 | -258.6727 | 1 | 0 | -192 | 0 | 1 | -258.6727
11 | ACAL | standard_control_drift | 20 | 1800.1 | 0.29139 | -179.12 | 11 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -174.6646 | 3 | 0 | -179.12 | -179.12 | 3897 | 3897 | 933 | 7189 | Normal | 3897 | 18.23 | -179.12 | -188.0043 | -20.0087 | 174.2799 | -187.9691 | 2 | 1 | 25.199 | 1.0633 | -187.9691 | -174.6646 | 1 | 0 | -192 | 0 | 1 | -174.6646
11 | ACAL | standard_control_drift | 20 | 1940.4 | 0.3203 | -97.684 | 11 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.6481 | 3 | 0 | -97.684 | -97.684 | 3997 | 3997 | 933 | 7189 | Normal | 3997 | 19.685 | -97.684 | -106.0932 | -14.6433 | 128.9481 | -105.1908 | 2 | 1 | 25.199 | 1.0633 | -105.1908 | -86.6481 | 1 | 0 | -192 | 0 | 1 | -86.6481
12 | ACAL | standard_control_drift | 20 | 1650.3 | 0.26349 | -257.505 | 11 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -258.6339 | 3 | 0 | -257.505 | -257.505 | 3950 | 3950 | 933 | 7189 | Normal | 3950 | 17.712 | -257.505 | -266.136 | -25.1264 | 217.5201 | -266.9412 | 2 | 1 | 25.199 | 1.0633 | -266.9412 | -258.6339 | 1 | 0 | -192 | 0 | 1 | -258.6339
12 | ACAL | standard_control_drift | 20 | 1800.3 | 0.29131 | -179.131 | 11 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -174.6378 | 3 | 0 | -179.131 | -179.131 | 3904 | 3904 | 933 | 7189 | Normal | 3904 | 18.302 | -179.131 | -187.9817 | -20.0072 | 174.2674 | -187.9439 | 2 | 1 | 25.199 | 1.0633 | -187.9439 | -174.6378 | 1 | 0 | -192 | 0 | 1 | -174.6378
12 | ACAL | standard_control_drift | 20 | 1940.8 | 0.32017 | -97.783 | 11 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.7229 | 3 | 0 | -97.783 | -97.783 | 4005 | 4005 | 933 | 7189 | Normal | 4005 | 19.774 | -97.783 | -106.1547 | -14.6473 | 128.9821 | -105.2611 | 2 | 1 | 25.199 | 1.0633 | -105.2611 | -86.7229 | 1 | 0 | -192 | 0 | 1 | -86.7229
13 | LCA | size | 30 | 1316.3 | 0.29195 | -177.392 | 13 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -163.8208 | 3 | 0 | -177.392 | -177.392 | 5999 | 5999 | 933 | 7189 | Normal | 5999 | 28.19 | -177.392 | -178.1876 | -19.3656 | 168.847 | -177.7706 | 2 | 1 | 25.199 | 1.0633 | -177.7706 | -163.8208 | 1 | 0 | -192 | 0 | 1 | -163.8208
13 | LCA | size | 30 | 1489.5 | 0.32457 | -85.482 | 13 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -67.3266 | 3 | 0 | -85.482 | -85.482 | 6162 | 6162 | 933 | 7189 | Normal | 6162 | 29.804 | -85.482 | -85.7749 | -13.3124 | 117.7034 | -87.0192 | 2 | 1 | 25.199 | 1.0633 | -87.0192 | -67.3266 | 1 | 0 | -192 | 0 | 1 | -67.3266
13 | LCA | size | 30 | 1650.7 | 0.279 | -213.878 | 13 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.2365 | 3 | 0 | -213.878 | -213.878 | 6198 | 6198 | 933 | 7189 | Normal | 6198 | 31.217 | -213.878 | -214.0617 | -21.7155 | 188.7007 | -212.9597 | 2 | 1 | 25.199 | 1.0633 | -212.9597 | -201.2365 | 1 | 0 | -192 | 0 | 1 | -201.2365
13 | LCA | size | 30 | 1800.7 | 0.27889 | -214.19 | 13 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -201.6393 | 3 | 0 | -214.19 | -214.19 | 6177 | 6177 | 933 | 7189 | Normal | 6177 | 32.322 | -214.19 | -214.4373 | -21.7401 | 188.9086 | -213.3385 | 2 | 1 | 25.199 | 1.0633 | -213.3385 | -201.6393 | 1 | 0 | -192 | 0 | 1 | -201.6393
13 | LCA | size | 30 | 1941.2 | 0.28487 | -197.354 | 13 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -183.9023 | 3 | 0 | -197.354 | -197.354 | 6224 | 6224 | 933 | 7189 | Normal | 6224 | 33.99 | -197.354 | -197.4592 | -20.628 | 179.5125 | -196.6571 | 2 | 1 | 25.199 | 1.0633 | -196.6571 | -183.9023 | 1 | 0 | -192 | 0 | 1 | -183.9023
13 | LCA | size | 30 | 2074.1 | 0.3044 | -142.314 | 13 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -124.2671 | 3 | 0 | -142.314 | -142.314 | 7060 | 7060 | 933 | 7189 | Normal | 7060 | 42.321 | -142.314 | -140.056 | -16.8679 | 147.744 | -140.571 | 2 | 1 | 25.199 | 1.0633 | -140.571 | -124.2671 | 1 | 0 | -192 | 0 | 1 | -124.2671
13 | LCA | size | 30 | 2217.9 | 0.3023 | -148.242 | 13 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -133.2496 | 3 | 0 | -148.242 | -148.242 | 6027 | 6027 | 933 | 7189 | Normal | 6027 | 41.385 | -148.242 | -148.9503 | -17.4505 | 152.6663 | -149.0188 | 2 | 1 | 25.199 | 1.0633 | -149.0188 | -133.2496 | 1 | 0 | -192 | 0 | 1 | -133.2496
13 | LCA | size | 30 | 2396.4 | 0.28379 | -200.388 | 13 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -192.7088 | 3 | 0 | -200.388 | -200.388 | 4822 | 4822 | 933 | 7189 | Normal | 4822 | 39.688 | -200.388 | -205.2788 | -21.1402 | 183.84 | -204.9395 | 2 | 1 | 25.199 | 1.0633 | -204.9395 | -192.7088 | 1 | 0 | -192 | 0 | 1 | -192.7088
14 | LCA | size | 25 | 1317.7 | 0.29297 | -174.285 | 13 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -164.9396 | 3 | 0 | -174.285 | -174.285 | 4839 | 4839 | 933 | 7189 | Normal | 4839 | 20.81 | -174.285 | -179.1098 | -19.426 | 169.3574 | -178.8229 | 2 | 1 | 25.199 | 1.0633 | -178.8229 | -164.9396 | 1 | 0 | -192 | 0 | 1 | -164.9396
14 | LCA | size | 25 | 1491 | 0.32596 | -81.312 | 13 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -66.0615 | 3 | 0 | -81.312 | -81.312 | 4926 | 4926 | 933 | 7189 | Normal | 4926 | 22.182 | -81.312 | -85.8027 | -13.3142 | 117.7187 | -85.8294 | 2 | 1 | 25.199 | 1.0633 | -85.8294 | -66.0615 | 1 | 0 | -192 | 0 | 1 | -66.0615
14 | LCA | size | 25 | 1652.1 | 0.28049 | -209.482 | 13 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.7778 | 3 | 0 | -209.482 | -209.482 | 4945 | 4945 | 933 | 7189 | Normal | 4945 | 23.358 | -209.482 | -213.9005 | -21.7049 | 188.6115 | -213.4688 | 2 | 1 | 25.199 | 1.0633 | -213.4688 | -201.7778 | 1 | 0 | -192 | 0 | 1 | -201.7778
14 | LCA | size | 25 | 1802.2 | 0.28009 | -210.586 | 13 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -202.9893 | 3 | 0 | -210.586 | -210.586 | 4936 | 4936 | 933 | 7189 | Normal | 4936 | 24.264 | -210.586 | -215.0386 | -21.7795 | 189.2414 | -214.6082 | 2 | 1 | 25.199 | 1.0633 | -214.6082 | -202.9893 | 1 | 0 | -192 | 0 | 1 | -202.9893
14 | LCA | size | 25 | 1942.4 | 0.28659 | -192.284 | 13 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -183.3669 | 3 | 0 | -192.284 | -192.284 | 4989 | 4989 | 933 | 7189 | Normal | 4989 | 25.654 | -192.284 | -196.5364 | -20.5675 | 179.0018 | -196.1535 | 2 | 1 | 25.199 | 1.0633 | -196.1535 | -183.3669 | 1 | 0 | -192 | 0 | 1 | -183.3669
14 | LCA | size | 25 | 2075.4 | 0.30576 | -138.248 | 13 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -123.773 | 3 | 0 | -138.248 | -138.248 | 5703 | 5703 | 933 | 7189 | Normal | 5703 | 32.363 | -138.248 | -139.9924 | -16.8638 | 147.7088 | -140.1063 | 2 | 1 | 25.199 | 1.0633 | -140.1063 | -123.773 | 1 | 0 | -192 | 0 | 1 | -123.773
14 | LCA | size | 25 | 2219 | 0.30325 | -145.324 | 13 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -133.789 | 3 | 0 | -145.324 | -145.324 | 4952 | 4952 | 933 | 7189 | Normal | 4952 | 31.755 | -145.324 | -149.716 | -17.5007 | 153.0901 | -149.5262 | 2 | 1 | 25.199 | 1.0633 | -149.5262 | -133.789 | 1 | 0 | -192 | 0 | 1 | -133.789
14 | LCA | size | 25 | 2397 | 0.28518 | -196.259 | 13 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -192.445 | 3 | 0 | -196.259 | -196.259 | 4012 | 4012 | 933 | 7189 | Normal | 4012 | 30.735 | -196.259 | -204.598 | -21.0956 | 183.4633 | -204.6914 | 2 | 1 | 25.199 | 1.0633 | -204.6914 | -192.445 | 1 | 0 | -192 | 0 | 1 | -192.445
15 | LCA | size | 20 | 1314.6 | 0.29442 | -170.354 | 13 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -164.0643 | 3 | 0 | -170.354 | -170.354 | 4122 | 4122 | 933 | 7189 | Normal | 4122 | 17.181 | -170.354 | -178.1858 | -19.3655 | 168.8461 | -177.9997 | 2 | 1 | 25.199 | 1.0633 | -177.9997 | -164.0643 | 1 | 0 | -192 | 0 | 1 | -164.0643
15 | LCA | size | 20 | 1487.7 | 0.32624 | -80.712 | 13 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -67.6778 | 3 | 0 | -80.712 | -80.712 | 4206 | 4206 | 933 | 7189 | Normal | 4206 | 18.349 | -80.712 | -88.1655 | -13.469 | 119.0264 | -87.3494 | 2 | 1 | 25.199 | 1.0633 | -87.3494 | -67.6778 | 1 | 0 | -192 | 0 | 1 | -67.6778
15 | LCA | size | 20 | 1648.8 | 0.28183 | -205.85 | 13 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.4007 | 3 | 0 | -205.85 | -205.85 | 4251 | 4251 | 933 | 7189 | Normal | 4251 | 19.283 | -205.85 | -213.104 | -21.6527 | 188.1707 | -213.1141 | 2 | 1 | 25.199 | 1.0633 | -213.1141 | -201.4007 | 1 | 0 | -192 | 0 | 1 | -201.4007
15 | LCA | size | 20 | 1798.9 | 0.281 | -208.177 | 13 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -204.0775 | 3 | 0 | -208.177 | -208.177 | 4217 | 4217 | 933 | 7189 | Normal | 4217 | 20.033 | -208.177 | -215.5816 | -21.815 | 189.5419 | -215.6316 | 2 | 1 | 25.199 | 1.0633 | -215.6316 | -204.0775 | 1 | 0 | -192 | 0 | 1 | -204.0775
15 | LCA | size | 20 | 1939.3 | 0.28734 | -190.304 | 13 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -184.6966 | 3 | 0 | -190.304 | -190.304 | 4265 | 4265 | 933 | 7189 | Normal | 4265 | 21.148 | -190.304 | -197.4963 | -20.6304 | 179.533 | -197.4041 | 2 | 1 | 25.199 | 1.0633 | -197.4041 | -184.6966 | 1 | 0 | -192 | 0 | 1 | -184.6966
15 | LCA | size | 20 | 2072.2 | 0.30682 | -135.425 | 13 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -123.5217 | 3 | 0 | -135.425 | -135.425 | 4892 | 4892 | 933 | 7189 | Normal | 4892 | 26.689 | -135.425 | -140.0455 | -16.8673 | 147.7382 | -139.8699 | 2 | 1 | 25.199 | 1.0633 | -139.8699 | -123.5217 | 1 | 0 | -192 | 0 | 1 | -123.5217
15 | LCA | size | 20 | 2215.6 | 0.30453 | -141.891 | 13 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.9797 | 3 | 0 | -141.891 | -141.891 | 4250 | 4250 | 933 | 7189 | Normal | 4250 | 26.192 | -141.891 | -149.1494 | -17.4636 | 152.7765 | -148.7651 | 2 | 1 | 25.199 | 1.0633 | -148.7651 | -132.9797 | 1 | 0 | -192 | 0 | 1 | -132.9797
15 | LCA | size | 20 | 2393.3 | 0.28607 | -193.884 | 13 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -193.0576 | 3 | 0 | -193.884 | -193.884 | 3489 | 3489 | 933 | 7189 | Normal | 3489 | 25.292 | -193.884 | -204.842 | -21.1116 | 183.5983 | -205.2676 | 2 | 1 | 25.199 | 1.0633 | -205.2676 | -193.0576 | 1 | 0 | -192 | 0 | 1 | -193.0576
16 | LCA | size | 15 | 1314.4 | 0.29591 | -165.9 | 13 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -163.6922 | 3 | 0 | -165.9 | -165.9 | 3333 | 3333 | 933 | 7189 | Normal | 3333 | 13.207 | -165.9 | -177.7157 | -19.3347 | 168.5859 | -177.6498 | 2 | 1 | 25.199 | 1.0633 | -177.6498 | -163.6922 | 1 | 0 | -192 | 0 | 1 | -163.6922
16 | LCA | size | 15 | 1487.5 | 0.32739 | -77.159 | 13 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -66.8697 | 3 | 0 | -77.159 | -77.159 | 3420 | 3420 | 933 | 7189 | Normal | 3420 | 14.284 | -77.159 | -88.4915 | -13.4904 | 119.2068 | -86.5894 | 2 | 1 | 25.199 | 1.0633 | -86.5894 | -66.8697 | 1 | 0 | -192 | 0 | 1 | -66.8697
16 | LCA | size | 15 | 1648.6 | 0.28318 | -201.788 | 13 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.8108 | 3 | 0 | -201.788 | -201.788 | 3461 | 3461 | 933 | 7189 | Normal | 3461 | 15.052 | -201.788 | -212.8971 | -21.6392 | 188.0562 | -213.4998 | 2 | 1 | 25.199 | 1.0633 | -213.4998 | -201.8108 | 1 | 0 | -192 | 0 | 1 | -201.8108
16 | LCA | size | 15 | 1798.9 | 0.28247 | -203.79 | 13 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -204.5534 | 3 | 0 | -203.79 | -203.79 | 3378 | 3378 | 933 | 7189 | Normal | 3378 | 15.557 | -203.79 | -215.3542 | -21.8001 | 189.4161 | -216.0792 | 2 | 1 | 25.199 | 1.0633 | -216.0792 | -204.5534 | 1 | 0 | -192 | 0 | 1 | -204.5534
16 | LCA | size | 15 | 1939.1 | 0.28931 | -184.504 | 13 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -183.5198 | 3 | 0 | -184.504 | -184.504 | 3392 | 3392 | 933 | 7189 | Normal | 3392 | 16.183 | -184.504 | -195.9907 | -20.5318 | 178.6997 | -196.2974 | 2 | 1 | 25.199 | 1.0633 | -196.2974 | -183.5198 | 1 | 0 | -192 | 0 | 1 | -183.5198
16 | LCA | size | 15 | 2071.8 | 0.30797 | -131.904 | 13 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -124.2584 | 3 | 0 | -131.904 | -131.904 | 3814 | 3814 | 933 | 7189 | Normal | 3814 | 19.871 | -131.904 | -141.192 | -16.9423 | 148.3727 | -140.5627 | 2 | 1 | 25.199 | 1.0633 | -140.5627 | -124.2584 | 1 | 0 | -192 | 0 | 1 | -124.2584
16 | LCA | size | 15 | 2215 | 0.3062 | -136.915 | 13 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.2491 | 3 | 0 | -136.915 | -136.915 | 3323 | 3323 | 933 | 7189 | Normal | 3323 | 18.915 | -136.915 | -148.787 | -17.4398 | 152.576 | -148.0779 | 2 | 1 | 25.199 | 1.0633 | -148.0779 | -132.2491 | 1 | 0 | -192 | 0 | 1 | -132.2491
16 | LCA | size | 15 | 2392.2 | 0.28854 | -186.67 | 13 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -191.3157 | 3 | 0 | -186.67 | -186.67 | 2669 | 2669 | 933 | 7189 | Normal | 2669 | 17.632 | -186.67 | -202.6515 | -20.9681 | 182.386 | -203.6293 | 2 | 1 | 25.199 | 1.0633 | -203.6293 | -191.3157 | 1 | 0 | -192 | 0 | 1 | -191.3157
17 | LCA | size | 10 | 1314.6 | 0.29799 | -160.036 | 13 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -164.5739 | 3 | 0 | -160.036 | -160.036 | 2361 | 2361 | 933 | 7189 | Normal | 2361 | 9.733 | -160.036 | -178.3166 | -19.3741 | 168.9185 | -178.479 | 2 | 1 | 25.199 | 1.0633 | -178.479 | -164.5739 | 1 | 0 | -192 | 0 | 1 | -164.5739
17 | LCA | size | 10 | 1488.1 | 0.32912 | -72.283 | 13 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -66.3581 | 3 | 0 | -72.283 | -72.283 | 2478 | 2478 | 933 | 7189 | Normal | 2478 | 10.288 | -72.283 | -89.6567 | -13.5667 | 119.8517 | -86.1082 | 2 | 1 | 25.199 | 1.0633 | -86.1082 | -66.3581 | 1 | 0 | -192 | 0 | 1 | -66.3581
17 | LCA | size | 10 | 1649.2 | 0.28577 | -194.484 | 13 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.2024 | 3 | 0 | -194.484 | -194.484 | 2529 | 2529 | 933 | 7189 | Normal | 2529 | 10.823 | -194.484 | -211.4757 | -21.5461 | 187.2696 | -212.9276 | 2 | 1 | 25.199 | 1.0633 | -212.9276 | -201.2024 | 1 | 0 | -192 | 0 | 1 | -201.2024
17 | LCA | size | 10 | 1799.5 | 0.28494 | -196.825 | 13 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -203.801 | 3 | 0 | -196.825 | -196.825 | 2528 | 2528 | 933 | 7189 | Normal | 2528 | 11.221 | -196.825 | -213.8242 | -21.6999 | 188.5693 | -215.3716 | 2 | 1 | 25.199 | 1.0633 | -215.3716 | -203.801 | 1 | 0 | -192 | 0 | 1 | -203.801
17 | LCA | size | 10 | 1939.7 | 0.29069 | -180.619 | 13 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -185.7154 | 3 | 0 | -180.619 | -180.619 | 2546 | 2546 | 933 | 7189 | Normal | 2546 | 11.838 | -180.619 | -197.4851 | -20.6297 | 179.5268 | -198.3623 | 2 | 1 | 25.199 | 1.0633 | -198.3623 | -185.7154 | 1 | 0 | -192 | 0 | 1 | -185.7154
17 | LCA | size | 10 | 2072.4 | 0.30978 | -126.781 | 13 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -123.2533 | 3 | 0 | -126.781 | -126.781 | 2978 | 2978 | 933 | 7189 | Normal | 2978 | 14.867 | -126.781 | -140.7084 | -16.9107 | 148.105 | -139.6175 | 2 | 1 | 25.199 | 1.0633 | -139.6175 | -123.2533 | 1 | 0 | -192 | 0 | 1 | -123.2533
17 | LCA | size | 10 | 2215.4 | 0.30742 | -133.441 | 13 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.7312 | 3 | 0 | -133.441 | -133.441 | 2657 | 2657 | 933 | 7189 | Normal | 2657 | 14.481 | -133.441 | -149.507 | -17.487 | 152.9744 | -148.5313 | 2 | 1 | 25.199 | 1.0633 | -148.5313 | -132.7312 | 1 | 0 | -192 | 0 | 1 | -132.7312
17 | LCA | size | 10 | 2392.4 | 0.29067 | -180.675 | 13 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -189.0102 | 3 | 0 | -180.675 | -180.675 | 2211 | 2211 | 933 | 7189 | Normal | 2211 | 13.894 | -180.675 | -200.1864 | -20.8066 | 181.0218 | -201.461 | 2 | 1 | 25.199 | 1.0633 | -201.461 | -189.0102 | 1 | 0 | -192 | 0 | 1 | -189.0102
18 | LCA | size | 5 | 1314.4 | 0.30418 | -142.479 | 13 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -163.0465 | 3 | 0 | -142.479 | -142.479 | 1018 | 1018 | 933 | 7189 | Normal | 1018 | 3.703 | -142.479 | -176.5335 | -19.2573 | 167.9316 | -177.0424 | 2 | 1 | 25.199 | 1.0633 | -177.0424 | -163.0465 | 1 | 0 | -192 | 0 | 1 | -163.0465
18 | LCA | size | 5 | 1487.5 | 0.33356 | -59.665 | 13 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -65.5326 | 3 | 0 | -59.665 | -59.665 | 1060 | 1060 | 933 | 7189 | Normal | 1060 | 4.049 | -59.665 | -92.9614 | -13.7831 | 121.6806 | -85.3319 | 2 | 1 | 25.199 | 1.0633 | -85.3319 | -65.5326 | 1 | 0 | -192 | 0 | 1 | -65.5326
18 | LCA | size | 5 | 1648.6 | 0.29197 | -176.916 | 13 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.8739 | 3 | 0 | -176.916 | -176.916 | 1079 | 1079 | 933 | 7189 | Normal | 1079 | 4.308 | -176.916 | -209.8793 | -21.4415 | 186.3861 | -213.5591 | 2 | 1 | 25.199 | 1.0633 | -213.5591 | -201.8739 | 1 | 0 | -192 | 0 | 1 | -201.8739
18 | LCA | size | 5 | 1798.4 | 0.29141 | -178.5 | 13 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -204.0441 | 3 | 0 | -178.5 | -178.5 | 1065 | 1065 | 933 | 7189 | Normal | 1065 | 4.49 | -178.5 | -211.7082 | -21.5613 | 187.3982 | -215.6002 | 2 | 1 | 25.199 | 1.0633 | -215.6002 | -204.0441 | 1 | 0 | -192 | 0 | 1 | -204.0441
18 | LCA | size | 5 | 1938.9 | 0.29759 | -161.063 | 13 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -183.6641 | 3 | 0 | -161.063 | -161.063 | 1067 | 1067 | 933 | 7189 | Normal | 1067 | 4.708 | -161.063 | -194.236 | -20.4168 | 177.7287 | -196.4331 | 2 | 1 | 25.199 | 1.0633 | -196.4331 | -183.6641 | 1 | 0 | -192 | 0 | 1 | -183.6641
18 | LCA | size | 5 | 2071.2 | 0.31422 | -114.182 | 13 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -125.804 | 3 | 0 | -114.182 | -114.182 | 1250 | 1250 | 933 | 7189 | Normal | 1250 | 5.835 | -114.182 | -144.3869 | -17.1516 | 150.1408 | -142.0164 | 2 | 1 | 25.199 | 1.0633 | -142.0164 | -125.804 | 1 | 0 | -192 | 0 | 1 | -125.804
18 | LCA | size | 5 | 2213.9 | 0.31335 | -116.645 | 13 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -131.1323 | 3 | 0 | -116.645 | -116.645 | 1106 | 1106 | 933 | 7189 | Normal | 1106 | 5.581 | -116.645 | -149.1449 | -17.4633 | 152.774 | -147.0276 | 2 | 1 | 25.199 | 1.0633 | -147.0276 | -131.1323 | 1 | 0 | -192 | 0 | 1 | -131.1323
19 | C4 | control | 20 | 939 | 0.30763 | -132.394 | 19 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -119.603 | 3 | 0 | -132.394 | -132.394 | 5092 | 5092 | 933 | 7189 | Normal | 5092 | 19.907 | -132.394 | -136.2632 | -16.6195 | 145.645 | -136.1844 | 2 | 1 | 25.199 | 1.0633 | -136.1844 | -119.603 | 1 | 0 | -192 | 0 | 1 | -119.603
19 | C4 | control | 20 | 1130.9 | 0.32977 | -69.977 | 19 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -53.0088 | 3 | 0 | -69.977 | -69.977 | 5331 | 5331 | 933 | 7189 | Normal | 5331 | 23.016 | -69.977 | -72.9862 | -12.4747 | 110.6257 | -73.5534 | 2 | 1 | 25.199 | 1.0633 | -73.5534 | -53.0088 | 1 | 0 | -192 | 0 | 1 | -53.0088
19 | C4 | control | 20 | 1316.1 | 0.2871 | -190.315 | 19 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -178.704 | 3 | 0 | -190.315 | -190.315 | 5625 | 5625 | 933 | 7189 | Normal | 5625 | 25.378 | -190.315 | -192.3176 | -20.2912 | 176.667 | -191.7681 | 2 | 1 | 25.199 | 1.0633 | -191.7681 | -178.704 | 1 | 0 | -192 | 0 | 1 | -178.704
19 | C4 | control | 20 | 1489.3 | 0.32483 | -83.904 | 19 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -66.6107 | 3 | 0 | -83.904 | -83.904 | 5759 | 5759 | 933 | 7189 | Normal | 5759 | 27.385 | -83.904 | -85.4651 | -13.2921 | 117.532 | -86.3458 | 2 | 1 | 25.199 | 1.0633 | -86.3458 | -66.6107 | 1 | 0 | -192 | 0 | 1 | -66.6107
19 | C4 | control | 20 | 1650.5 | 0.26146 | -262.617 | 19 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -253.8018 | 3 | 0 | -262.617 | -262.617 | 5808 | 5808 | 933 | 7189 | Normal | 5808 | 28.734 | -262.617 | -264.0193 | -24.9878 | 216.3486 | -262.3967 | 2 | 1 | 25.199 | 1.0633 | -262.3967 | -253.8018 | 1 | 0 | -192 | 0 | 1 | -253.8018
20 | C4 | control | 20 | 940.3 | 0.30787 | -132.187 | 19 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -119.8885 | 3 | 0 | -132.187 | -132.187 | 4948 | 4948 | 933 | 7189 | Normal | 4948 | 18.82 | -132.187 | -136.5941 | -16.6412 | 145.8281 | -136.4529 | 2 | 1 | 25.199 | 1.0633 | -136.4529 | -119.8885 | 1 | 0 | -192 | 0 | 1 | -119.8885
20 | C4 | control | 20 | 1131.9 | 0.32982 | -70.319 | 19 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -53.6056 | 3 | 0 | -70.319 | -70.319 | 5236 | 5236 | 933 | 7189 | Normal | 5236 | 21.904 | -70.319 | -73.6653 | -12.5192 | 111.0016 | -74.1147 | 2 | 1 | 25.199 | 1.0633 | -74.1147 | -53.6056 | 1 | 0 | -192 | 0 | 1 | -53.6056
20 | C4 | control | 20 | 1316.9 | 0.28742 | -189.839 | 19 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -178.6136 | 3 | 0 | -189.839 | -189.839 | 5518 | 5518 | 933 | 7189 | Normal | 5518 | 24.645 | -189.839 | -192.2017 | -20.2836 | 176.6028 | -191.6831 | 2 | 1 | 25.199 | 1.0633 | -191.6831 | -178.6136 | 1 | 0 | -192 | 0 | 1 | -178.6136
20 | C4 | control | 20 | 1490.2 | 0.32518 | -83.405 | 19 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -66.2178 | 3 | 0 | -83.405 | -83.405 | 5705 | 5705 | 933 | 7189 | Normal | 5705 | 26.823 | -83.405 | -85.1428 | -13.271 | 117.3536 | -85.9763 | 2 | 1 | 25.199 | 1.0633 | -85.9763 | -66.2178 | 1 | 0 | -192 | 0 | 1 | -66.2178
20 | C4 | control | 20 | 1651.3 | 0.26146 | -263.023 | 19 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -254.5053 | 3 | 0 | -263.023 | -263.023 | 5747 | 5747 | 933 | 7189 | Normal | 5747 | 28.245 | -263.023 | -264.6233 | -25.0273 | 216.6829 | -263.0583 | 2 | 1 | 25.199 | 1.0633 | -263.0583 | -254.5053 | 1 | 0 | -192 | 0 | 1 | -254.5053
21 | ACAL | standard_control_drift | 20 | 1649.6 | 0.2637 | -256.715 | 21 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -257.3196 | 3 | 0 | -256.715 | -256.715 | 4019 | 4019 | 933 | 7189 | Normal | 4019 | 18.155 | -256.715 | -265.0213 | -25.0534 | 216.9032 | -265.7051 | 3 | 1 | 25.199 | 1.0633 | -265.7051 | -257.3196 | 1 | 0 | -192 | 0 | 1 | -257.3196
21 | ACAL | standard_control_drift | 20 | 1799.7 | 0.2914 | -178.627 | 21 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -173.7037 | 3 | 0 | -178.627 | -178.627 | 3976 | 3976 | 933 | 7189 | Normal | 3976 | 18.687 | -178.627 | -187.135 | -19.9517 | 173.7988 | -187.0654 | 3 | 1 | 25.199 | 1.0633 | -187.0654 | -173.7037 | 1 | 0 | -192 | 0 | 1 | -173.7037
21 | ACAL | standard_control_drift | 20 | 1940.1 | 0.32004 | -97.92 | 21 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.5892 | 3 | 0 | -97.92 | -97.92 | 4077 | 4077 | 933 | 7189 | Normal | 4077 | 20.196 | -97.92 | -105.9576 | -14.6344 | 128.8731 | -105.1354 | 3 | 1 | 25.199 | 1.0633 | -105.1354 | -86.5892 | 1 | 0 | -192 | 0 | 1 | -86.5892
22 | ACAL | standard_control_drift | 20 | 1649.8 | 0.26431 | -255.148 | 21 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -256.7389 | 3 | 0 | -255.148 | -255.148 | 3857 | 3857 | 933 | 7189 | Normal | 3857 | 17.284 | -255.148 | -264.2258 | -25.0013 | 216.4629 | -265.1589 | 3 | 1 | 25.199 | 1.0633 | -265.1589 | -256.7389 | 1 | 0 | -192 | 0 | 1 | -256.7389
22 | ACAL | standard_control_drift | 20 | 1799.9 | 0.29197 | -177.19 | 21 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -173.241 | 3 | 0 | -177.19 | -177.19 | 3780 | 3780 | 933 | 7189 | Normal | 3780 | 17.615 | -177.19 | -186.6459 | -19.9197 | 173.5281 | -186.6302 | 3 | 1 | 25.199 | 1.0633 | -186.6302 | -173.241 | 1 | 0 | -192 | 0 | 1 | -173.241
22 | ACAL | standard_control_drift | 20 | 1940.1 | 0.32052 | -96.743 | 21 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.3722 | 3 | 0 | -96.743 | -96.743 | 3815 | 3815 | 933 | 7189 | Normal | 3815 | 18.619 | -96.743 | -106.0261 | -14.6389 | 128.9109 | -104.9313 | 3 | 1 | 25.199 | 1.0633 | -104.9313 | -86.3722 | 1 | 0 | -192 | 0 | 1 | -86.3722
23 | ACAL | control | 10 | 1487 | 0.33101 | -66.851 | 21 | C23 Alkane | C23 | Alkane | ACAL | 1 | -67.2 |  |  | -63.9184 | 3 | 0 | -66.851 | -66.851 | 1958 | 1958 | 933 | 7189 | Normal | 1958 | 7.645 | -66.851 | -88.641 | -13.5002 | 119.2896 | -83.8137 | 3 | 1 | 25.199 | 1.0633 | -83.8137 | -63.9184 | 1 | 0 | -192 | 0 | 1 | -63.9184
23 | ACAL | control | 10 | 1648.2 | 0.26969 | -239.719 | 21 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -257.9677 | 3 | 0 | -239.719 | -239.719 | 1974 | 1974 | 933 | 7189 | Normal | 1974 | 8.1 | -239.719 | -261.3564 | -24.8134 | 214.8749 | -266.3147 | 3 | 1 | 25.199 | 1.0633 | -266.3147 | -257.9677 | 1 | 0 | -192 | 0 | 1 | -257.9677
23 | ACAL | control | 10 | 1798.2 | 0.29567 | -166.488 | 21 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -176.2745 | 3 | 0 | -166.488 | -166.488 | 1925 | 1925 | 933 | 7189 | Normal | 1925 | 8.265 | -166.488 | -188.5967 | -20.0475 | 174.6078 | -189.4832 | 3 | 1 | 25.199 | 1.0633 | -189.4832 | -176.2745 | 1 | 0 | -192 | 0 | 1 | -176.2745
23 | ACAL | control | 10 | 1938.3 | 0.32343 | -88.227 | 21 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -88.0021 | 3 | 0 | -88.227 | -88.227 | 1951 | 1951 | 933 | 7189 | Normal | 1951 | 8.776 | -88.227 | -110.0842 | -14.9047 | 131.1568 | -106.4642 | 3 | 1 | 25.199 | 1.0633 | -106.4642 | -88.0021 | 1 | 0 | -192 | 0 | 1 | -88.0021
23 | ACAL | control | 10 | 2070.6 | 0.26754 | -245.793 | 21 | C31 Alkane | C31 | Alkane | ACAL | 1 | -271.9 |  |  | -266.8215 | 3 | 0 | -245.793 | -245.793 | 1836 | 1836 | 933 | 7189 | Normal | 1836 | 8.735 | -245.793 | -268.7893 | -25.3002 | 218.9885 | -274.6415 | 3 | 1 | 25.199 | 1.0633 | -274.6415 | -266.8215 | 1 | 0 | -192 | 0 | 1 | -266.8215
24 | ACAL | control | 10 | 1488.7 | 0.33085 | -67.513 | 21 | C23 Alkane | C23 | Alkane | ACAL | 1 | -67.2 |  |  | -64.734 | 3 | 0 | -67.513 | -67.513 | 1949 | 1949 | 933 | 7189 | Normal | 1949 | 7.684 | -67.513 | -89.3894 | -13.5492 | 119.7037 | -84.5808 | 3 | 1 | 25.199 | 1.0633 | -84.5808 | -64.734 | 1 | 0 | -192 | 0 | 1 | -64.734
24 | ACAL | control | 10 | 1649.8 | 0.27002 | -238.96 | 21 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -257.2127 | 3 | 0 | -238.96 | -238.96 | 1967 | 1967 | 933 | 7189 | Normal | 1967 | 8.049 | -238.96 | -260.664 | -24.768 | 214.4917 | -265.6046 | 3 | 1 | 25.199 | 1.0633 | -265.6046 | -257.2127 | 1 | 0 | -192 | 0 | 1 | -257.2127
24 | ACAL | control | 10 | 1799.7 | 0.29593 | -165.949 | 21 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -175.7031 | 3 | 0 | -165.949 | -165.949 | 1922 | 1922 | 933 | 7189 | Normal | 1922 | 8.238 | -165.949 | -188.087 | -20.0141 | 174.3256 | -188.9459 | 3 | 1 | 25.199 | 1.0633 | -188.9459 | -175.7031 | 1 | 0 | -192 | 0 | 1 | -175.7031
24 | ACAL | control | 10 | 1940.1 | 0.32312 | -89.324 | 21 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -89.3493 | 3 | 0 | -89.324 | -89.324 | 1938 | 1938 | 933 | 7189 | Normal | 1938 | 8.743 | -89.324 | -111.3065 | -14.9848 | 131.8333 | -107.7312 | 3 | 1 | 25.199 | 1.0633 | -107.7312 | -89.3493 | 1 | 0 | -192 | 0 | 1 | -89.3493
24 | ACAL | control | 10 | 2072.2 | 0.26716 | -247.03 | 21 | C31 Alkane | C31 | Alkane | ACAL | 1 | -271.9 |  |  | -268.4022 | 3 | 0 | -247.03 | -247.03 | 1824 | 1824 | 933 | 7189 | Normal | 1824 | 8.694 | -247.03 | -270.1493 | -25.3893 | 219.7411 | -276.1282 | 3 | 1 | 25.199 | 1.0633 | -276.1282 | -268.4022 | 1 | 0 | -192 | 0 | 1 | -268.4022
25 | ACAL | control | 5 | 1487 | 0.33335 | -60.021 | 21 | C23 Alkane | C23 | Alkane | ACAL | 1 | -67.2 |  |  | -67.277 | 3 | 0 | -60.021 | -60.021 | 979 | 979 | 933 | 7189 | Normal | 979 | 3.704 | -60.021 | -94.8079 | -13.9041 | 122.7025 | -86.9725 | 3 | 1 | 25.199 | 1.0633 | -86.9725 | -67.277 | 1 | 0 | -192 | 0 | 1 | -67.277
25 | ACAL | control | 5 | 1648.2 | 0.27583 | -222.223 | 21 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -257.0232 | 3 | 0 | -222.223 | -222.223 | 992 | 992 | 933 | 7189 | Normal | 992 | 3.941 | -222.223 | -256.7626 | -24.5125 | 212.3326 | -265.4264 | 3 | 1 | 25.199 | 1.0633 | -265.4264 | -257.0232 | 1 | 0 | -192 | 0 | 1 | -257.0232
25 | ACAL | control | 5 | 1798.2 | 0.30148 | -149.88 | 21 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -172.343 | 3 | 0 | -149.88 | -149.88 | 990 | 990 | 933 | 7189 | Normal | 990 | 4.108 | -149.88 | -184.4574 | -19.7763 | 172.317 | -185.7857 | 3 | 1 | 25.199 | 1.0633 | -185.7857 | -172.343 | 1 | 0 | -192 | 0 | 1 | -172.343
25 | ACAL | control | 5 | 1938.5 | 0.32624 | -80.078 | 21 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -90.0227 | 3 | 0 | -80.078 | -80.078 | 1021 | 1021 | 933 | 7189 | Normal | 1021 | 4.47 | -80.078 | -114.0773 | -15.1663 | 133.3667 | -108.3645 | 3 | 1 | 25.199 | 1.0633 | -108.3645 | -90.0227 | 1 | 0 | -192 | 0 | 1 | -90.0227
25 | ACAL | control | 5 | 2070.6 | 0.27391 | -227.629 | 21 | C31 Alkane | C31 | Alkane | ACAL | 1 | -271.9 |  |  | -263.2973 | 3 | 0 | -227.629 | -227.629 | 994 | 994 | 933 | 7189 | Normal | 994 | 4.621 | -227.629 | -262.1308 | -24.8641 | 215.3035 | -271.3271 | 3 | 1 | 25.199 | 1.0633 | -271.3271 | -263.2973 | 1 | 0 | -192 | 0 | 1 | -263.2973
26 | ACAL | control | 5 | 1487 | 0.33311 | -61.398 | 21 | C23 Alkane | C23 | Alkane | ACAL | 1 | -67.2 |  |  | -68.7705 | 3 | 0 | -61.398 | -61.398 | 986 | 986 | 933 | 7189 | Normal | 986 | 3.727 | -61.398 | -96.0513 | -13.9855 | 123.3906 | -88.3771 | 3 | 1 | 25.199 | 1.0633 | -88.3771 | -68.7705 | 1 | 0 | -192 | 0 | 1 | -68.7705
26 | ACAL | control | 5 | 1648.2 | 0.27655 | -220.78 | 21 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -255.305 | 3 | 0 | -220.78 | -220.78 | 993 | 993 | 933 | 7189 | Normal | 993 | 3.945 | -220.78 | -255.3007 | -24.4167 | 211.5235 | -263.8104 | 3 | 1 | 25.199 | 1.0633 | -263.8104 | -255.305 | 1 | 0 | -192 | 0 | 1 | -255.305
26 | ACAL | control | 5 | 1798 | 0.30061 | -152.961 | 21 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -176.1644 | 3 | 0 | -152.961 | -152.961 | 981 | 981 | 933 | 7189 | Normal | 981 | 4.075 | -152.961 | -187.7097 | -19.9894 | 174.1168 | -189.3797 | 3 | 1 | 25.199 | 1.0633 | -189.3797 | -176.1644 | 1 | 0 | -192 | 0 | 1 | -176.1644
26 | ACAL | control | 5 | 1938.5 | 0.32532 | -83.358 | 21 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -94.2393 | 3 | 0 | -83.358 | -83.358 | 1000 | 1000 | 933 | 7189 | Normal | 1000 | 4.385 | -83.358 | -117.747 | -15.4067 | 135.3976 | -112.3302 | 3 | 1 | 25.199 | 1.0633 | -112.3302 | -94.2393 | 1 | 0 | -192 | 0 | 1 | -94.2393
26 | ACAL | control | 5 | 2070.6 | 0.2743 | -227.116 | 21 | C31 Alkane | C31 | Alkane | ACAL | 1 | -271.9 |  |  | -263.5984 | 3 | 0 | -227.116 | -227.116 | 963 | 963 | 933 | 7189 | Normal | 963 | 4.462 | -227.116 | -262.2119 | -24.8694 | 215.3484 | -271.6103 | 3 | 1 | 25.199 | 1.0633 | -271.6103 | -263.5984 | 1 | 0 | -192 | 0 | 1 | -263.5984
27 | BOF1-11 | sample |  | 1483.5 | 0.28013 | -210.192 | 27 | C23 Alkane | C23 | Alkane | [5 empty] | -234.4265 | 3 | 0 | -210.192 | -210.192 | 1362 | 1362 | 933 | 7189 | Normal | 1362 | 5.271 | -210.192 | -238.7879 | -23.3351 | 202.3849 | -244.1744 | 3 | 1 | 25.199 | 1.0633 | -244.1744 | -234.4265 | 1 | 0 | -192 | 0 | 1 | -234.4265
27 | BOF1-11 | sample |  | 1649.4 | 0.26543 | -251.63 | 27 | C25 Alkane | C25 | Alkane | [5 empty] | -241.6493 | 3 | 0 | -251.63 | -251.63 | 5955 | 5955 | 933 | 7189 | Normal | 5955 | 29.805 | -251.63 | -252.5636 | -24.2374 | 210.0087 | -250.9674 | 3 | 1 | 25.199 | 1.0633 | -250.9674 | -241.6493 | 1 | 0 | -192 | 0 | 1 | -241.6493
27 | BOF1-11 | sample |  | 1802.6 | 0.26341 | -257.327 | 27 | C27 Alkane | C27 | Alkane | [5 empty] | -236.3462 | 3 | 0 | -257.327 | -257.327 | 9300 | 9300 | 933 | 7189 | Large | 9300 | 56.28 | -257.327 | -249.9019 | -24.0631 | 208.5357 | -245.9799 | 3 | 1 | 25.199 | 1.0633 | -245.9799 | -236.3462 | 1 | 0 | -192 | 0 | 1 | -236.3462
27 | BOF1-11 | sample |  | 1940.4 | 0.27391 | -227.744 | 27 | C29 Alkane | C29 | Alkane | [5 empty] | -226.4482 | 3 | 0 | -227.744 | -227.744 | 3981 | 3981 | 933 | 7189 | Normal | 3981 | 19.684 | -227.744 | -236.2284 | -23.1674 | 200.9684 | -236.6709 | 3 | 1 | 25.199 | 1.0633 | -236.6709 | -226.4482 | 1 | 0 | -192 | 0 | 1 | -226.4482
27 | BOF1-11 | sample |  | 2072.9 | 0.27519 | -224.135 | 27 | C31 Alkane | C31 | Alkane | [5 empty] | -225.4687 | 3 | 0 | -224.135 | -224.135 | 3541 | 3541 | 933 | 7189 | Normal | 3541 | 17.978 | -224.135 | -234.8156 | -23.0749 | 200.1865 | -235.7498 | 3 | 1 | 25.199 | 1.0633 | -235.7498 | -225.4687 | 1 | 0 | -192 | 0 | 1 | -225.4687
28 | BOF1-11 | sample |  | 1483.1 | 0.27924 | -212.613 | 27 | C23 Alkane | C23 | Alkane | [5 empty] | -236.956 | 3 | 0 | -212.613 | -212.613 | 1375 | 1375 | 933 | 7189 | Normal | 1375 | 5.323 | -212.613 | -241.0308 | -23.482 | 203.6262 | -246.5534 | 3 | 1 | 25.199 | 1.0633 | -246.5534 | -236.956 | 1 | 0 | -192 | 0 | 1 | -236.956
28 | BOF1-11 | sample |  | 1649.2 | 0.26522 | -252.168 | 27 | C25 Alkane | C25 | Alkane | [5 empty] | -242.2166 | 3 | 0 | -252.168 | -252.168 | 5954 | 5954 | 933 | 7189 | Normal | 5954 | 30.181 | -252.168 | -253.1048 | -24.2729 | 210.3082 | -251.5009 | 3 | 1 | 25.199 | 1.0633 | -251.5009 | -242.2166 | 1 | 0 | -192 | 0 | 1 | -242.2166
28 | BOF1-11 | sample |  | 1802.2 | 0.26346 | -257.123 | 27 | C27 Alkane | C27 | Alkane | [5 empty] | -235.8334 | 3 | 0 | -257.123 | -257.123 | 9416 | 9416 | 933 | 7189 | Large | 9416 | 57.292 | -257.123 | -249.4655 | -24.0345 | 208.2942 | -245.4977 | 3 | 1 | 25.199 | 1.0633 | -245.4977 | -235.8334 | 1 | 0 | -192 | 0 | 1 | -235.8334
28 | BOF1-11 | sample |  | 1940.1 | 0.27386 | -227.801 | 27 | C29 Alkane | C29 | Alkane | [5 empty] | -225.9973 | 3 | 0 | -227.801 | -227.801 | 4064 | 4064 | 933 | 7189 | Normal | 4064 | 20.34 | -227.801 | -235.8985 | -23.1458 | 200.7858 | -236.2469 | 3 | 1 | 25.199 | 1.0633 | -236.2469 | -225.9973 | 1 | 0 | -192 | 0 | 1 | -225.9973
28 | BOF1-11 | sample |  | 2072.4 | 0.27429 | -226.587 | 27 | C31 Alkane | C31 | Alkane | [5 empty] | -226.9505 | 3 | 0 | -226.587 | -226.587 | 3711 | 3711 | 933 | 7189 | Normal | 3711 | 18.929 | -226.587 | -236.3883 | -23.1779 | 201.0569 | -237.1434 | 3 | 1 | 25.199 | 1.0633 | -237.1434 | -226.9505 | 1 | 0 | -192 | 0 | 1 | -226.9505
29 | BOF1-31 | sample |  | 1484.1 | 0.26982 | -239.168 | 29 | C23 Alkane | C23 | Alkane | [5 empty] | -244.8547 | 3 | 0 | -239.168 | -239.168 | 3143 | 3143 | 933 | 7189 | Normal | 3143 | 13.154 | -239.168 | -252.0842 | -24.206 | 209.7434 | -253.982 | 3 | 1 | 25.199 | 1.0633 | -253.982 | -244.8547 | 1 | 0 | -192 | 0 | 1 | -244.8547
29 | BOF1-31 | sample |  | 1648.6 | 0.26652 | -248.482 | 29 | C25 Alkane | C25 | Alkane | [5 empty] | -239.347 | 3 | 0 | -248.482 | -248.482 | 5726 | 5726 | 933 | 7189 | Normal | 5726 | 28.615 | -248.482 | -250.1509 | -24.0794 | 208.6735 | -248.8021 | 3 | 1 | 25.199 | 1.0633 | -248.8021 | -239.347 | 1 | 0 | -192 | 0 | 1 | -239.347
29 | BOF1-31 | sample |  | 1800.1 | 0.26765 | -245.279 | 29 | C27 Alkane | C27 | Alkane | [5 empty] | -233.0074 | 3 | 0 | -245.279 | -245.279 | 6452 | 6452 | 933 | 7189 | Normal | 6452 | 34.326 | -245.279 | -244.7096 | -23.723 | 205.6621 | -242.8398 | 3 | 1 | 25.199 | 1.0633 | -242.8398 | -233.0074 | 1 | 0 | -192 | 0 | 1 | -233.0074
29 | BOF1-31 | sample |  | 1939.7 | 0.27112 | -235.514 | 29 | C29 Alkane | C29 | Alkane | [5 empty] | -233.1184 | 3 | 0 | -235.514 | -235.514 | 4254 | 4254 | 933 | 7189 | Normal | 4254 | 21.191 | -235.514 | -242.7548 | -23.5949 | 204.5803 | -242.9442 | 3 | 1 | 25.199 | 1.0633 | -242.9442 | -233.1184 | 1 | 0 | -192 | 0 | 1 | -233.1184
29 | BOF1-31 | sample |  | 2071.6 | 0.27255 | -231.464 | 29 | C31 Alkane | C31 | Alkane | [5 empty] | -235.9568 | 3 | 0 | -231.464 | -231.464 | 3204 | 3204 | 933 | 7189 | Normal | 3204 | 16.026 | -231.464 | -244.0198 | -23.6778 | 205.2804 | -245.6137 | 3 | 1 | 25.199 | 1.0633 | -245.6137 | -235.9568 | 1 | 0 | -192 | 0 | 1 | -235.9568
29 | BOF1-31 | sample |  | 2213.1 | 0.28081 | -208.194 | 29 | C33 Alkane | C33 | Alkane | [5 empty] | -239.0309 | 3 | 0 | -208.194 | -208.194 | 1051 | 1051 | 933 | 7189 | Normal | 1051 | 5.246 | -208.194 | -241.6503 | -23.5226 | 203.969 | -248.5048 | 3 | 1 | 25.199 | 1.0633 | -248.5048 | -239.0309 | 1 | 0 | -192 | 0 | 1 | -239.0309
30 | BOF1-31 | sample |  | 1484.7 | 0.26936 | -240.459 | 29 | C23 Alkane | C23 | Alkane | [5 empty] | -246.5242 | 3 | 0 | -240.459 | -240.459 | 3112 | 3112 | 933 | 7189 | Normal | 3112 | 13.015 | -240.459 | -253.5611 | -24.3028 | 210.5608 | -255.5522 | 3 | 1 | 25.199 | 1.0633 | -255.5522 | -246.5242 | 1 | 0 | -192 | 0 | 1 | -246.5242
30 | BOF1-31 | sample |  | 1649.2 | 0.26627 | -249.167 | 29 | C25 Alkane | C25 | Alkane | [5 empty] | -240.15 | 3 | 0 | -249.167 | -249.167 | 5707 | 5707 | 933 | 7189 | Normal | 5707 | 28.198 | -249.167 | -250.8982 | -24.1283 | 209.0871 | -249.5573 | 3 | 1 | 25.199 | 1.0633 | -249.5573 | -240.15 | 1 | 0 | -192 | 0 | 1 | -240.15
30 | BOF1-31 | sample |  | 1800.7 | 0.26778 | -244.905 | 29 | C27 Alkane | C27 | Alkane | [5 empty] | -233.2372 | 3 | 0 | -244.905 | -244.905 | 6293 | 6293 | 933 | 7189 | Normal | 6293 | 33.553 | -244.905 | -244.8035 | -23.7291 | 205.7141 | -243.056 | 3 | 1 | 25.199 | 1.0633 | -243.056 | -233.2372 | 1 | 0 | -192 | 0 | 1 | -233.2372
30 | BOF1-31 | sample |  | 1940.4 | 0.27098 | -235.896 | 29 | C29 Alkane | C29 | Alkane | [5 empty] | -234.238 | 3 | 0 | -235.896 | -235.896 | 4136 | 4136 | 933 | 7189 | Normal | 4136 | 20.41 | -235.896 | -243.6642 | -23.6545 | 205.0836 | -243.9972 | 3 | 1 | 25.199 | 1.0633 | -243.9972 | -234.238 | 1 | 0 | -192 | 0 | 1 | -234.238
30 | BOF1-31 | sample |  | 2072.2 | 0.27306 | -230.009 | 29 | C31 Alkane | C31 | Alkane | [5 empty] | -235.8541 | 3 | 0 | -230.009 | -230.009 | 3024 | 3024 | 933 | 7189 | Normal | 3024 | 15.101 | -230.009 | -243.649 | -23.6535 | 205.0751 | -245.5171 | 3 | 1 | 25.199 | 1.0633 | -245.5171 | -235.8541 | 1 | 0 | -192 | 0 | 1 | -235.8541
30 | BOF1-31 | sample |  | 2213.5 | 0.28137 | -206.58 | 29 | C33 Alkane | C33 | Alkane | [5 empty] | -239.1418 | 3 | 0 | -206.58 | -206.58 | 976 | 976 | 933 | 7189 | Normal | 976 | 4.822 | -206.58 | -241.4245 | -23.5078 | 203.8441 | -248.6091 | 3 | 1 | 25.199 | 1.0633 | -248.6091 | -239.1418 | 1 | 0 | -192 | 0 | 1 | -239.1418
31 | ACAL | standard_control_drift | 20 | 1650.1 | 0.26365 | -256.345 | 31 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -256.9023 | 3 | 0 | -256.345 | -256.345 | 4022 | 4022 | 933 | 7189 | Normal | 4022 | 17.977 | -256.345 | -264.6373 | -25.0283 | 216.6907 | -265.3127 | 4 | 1 | 25.199 | 1.0633 | -265.3127 | -256.9023 | 1 | 0 | -192 | 0 | 1 | -256.9023
31 | ACAL | standard_control_drift | 20 | 1800.1 | 0.29124 | -178.541 | 31 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -173.9372 | 3 | 0 | -178.541 | -178.541 | 3916 | 3916 | 933 | 7189 | Normal | 3916 | 18.284 | -178.541 | -187.3341 | -19.9648 | 173.909 | -187.2851 | 4 | 1 | 25.199 | 1.0633 | -187.2851 | -173.9372 | 1 | 0 | -192 | 0 | 1 | -173.9372
31 | ACAL | standard_control_drift | 20 | 1940.4 | 0.31967 | -98.344 | 31 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -87.5003 | 3 | 0 | -98.344 | -98.344 | 3961 | 3961 | 933 | 7189 | Normal | 3961 | 19.32 | -98.344 | -106.9229 | -14.6977 | 129.4072 | -105.9922 | 4 | 1 | 25.199 | 1.0633 | -105.9922 | -87.5003 | 1 | 0 | -192 | 0 | 1 | -87.5003
32 | ACAL | standard_control_drift | 20 | 1649.4 | 0.26409 | -255.241 | 31 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -255.991 | 3 | 0 | -255.241 | -255.241 | 3981 | 3981 | 933 | 7189 | Normal | 3981 | 17.959 | -255.241 | -263.7254 | -24.9685 | 216.186 | -264.4556 | 4 | 1 | 25.199 | 1.0633 | -264.4556 | -255.991 | 1 | 0 | -192 | 0 | 1 | -255.991
32 | ACAL | standard_control_drift | 20 | 1799.5 | 0.2912 | -178.791 | 31 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -174.0097 | 3 | 0 | -178.791 | -178.791 | 3952 | 3952 | 933 | 7189 | Normal | 3952 | 18.503 | -178.791 | -187.4125 | -19.9699 | 173.9524 | -187.3532 | 4 | 1 | 25.199 | 1.0633 | -187.3532 | -174.0097 | 1 | 0 | -192 | 0 | 1 | -174.0097
32 | ACAL | standard_control_drift | 20 | 1939.9 | 0.31985 | -97.993 | 31 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.7994 | 3 | 0 | -97.993 | -97.993 | 4043 | 4043 | 933 | 7189 | Normal | 4043 | 19.986 | -97.993 | -106.1877 | -14.6495 | 129.0004 | -105.3331 | 4 | 1 | 25.199 | 1.0633 | -105.3331 | -86.7994 | 1 | 0 | -192 | 0 | 1 | -86.7994
33 | C4 | control | 20 | 938.8 | 0.30794 | -131.91 | 33 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -118.5172 | 3 | 0 | -131.91 | -131.91 | 5262 | 5262 | 933 | 7189 | Normal | 5262 | 20.708 | -131.91 | -135.1634 | -16.5475 | 145.0363 | -135.1632 | 4 | 1 | 25.199 | 1.0633 | -135.1632 | -118.5172 | 1 | 0 | -192 | 0 | 1 | -118.5172
33 | C4 | control | 20 | 1130.9 | 0.32944 | -71.302 | 33 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -54.3808 | 3 | 0 | -71.302 | -71.302 | 5341 | 5341 | 933 | 7189 | Normal | 5341 | 23.15 | -71.302 | -74.276 | -12.5592 | 111.3396 | -74.8437 | 4 | 1 | 25.199 | 1.0633 | -74.8437 | -54.3808 | 1 | 0 | -192 | 0 | 1 | -54.3808
33 | C4 | control | 20 | 1316.1 | 0.28724 | -190.267 | 33 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -178.559 | 3 | 0 | -190.267 | -190.267 | 5650 | 5650 | 933 | 7189 | Normal | 5650 | 25.426 | -190.267 | -192.1864 | -20.2826 | 176.5944 | -191.6318 | 4 | 1 | 25.199 | 1.0633 | -191.6318 | -178.559 | 1 | 0 | -192 | 0 | 1 | -178.559
33 | C4 | control | 20 | 1489.1 | 0.32482 | -84.335 | 33 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -67.0049 | 3 | 0 | -84.335 | -84.335 | 5783 | 5783 | 933 | 7189 | Normal | 5783 | 27.473 | -84.335 | -85.8182 | -13.3152 | 117.7273 | -86.7166 | 4 | 1 | 25.199 | 1.0633 | -86.7166 | -67.0049 | 1 | 0 | -192 | 0 | 1 | -67.0049
33 | C4 | control | 20 | 1650.5 | 0.26154 | -262.727 | 33 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -253.7907 | 3 | 0 | -262.727 | -262.727 | 5836 | 5836 | 933 | 7189 | Normal | 5836 | 28.833 | -262.727 | -264.0391 | -24.9891 | 216.3596 | -262.3862 | 4 | 1 | 25.199 | 1.0633 | -262.3862 | -253.7907 | 1 | 0 | -192 | 0 | 1 | -253.7907
34 | C4 | control | 20 | 940.1 | 0.30793 | -131.66 | 33 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -118.2105 | 3 | 0 | -131.66 | -131.66 | 5275 | 5275 | 933 | 7189 | Normal | 5275 | 20.461 | -131.66 | -134.8672 | -16.5281 | 144.8724 | -134.8748 | 4 | 1 | 25.199 | 1.0633 | -134.8748 | -118.2105 | 1 | 0 | -192 | 0 | 1 | -118.2105
34 | C4 | control | 20 | 1131.7 | 0.3295 | -70.811 | 33 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -53.6199 | 3 | 0 | -70.811 | -70.811 | 5441 | 5441 | 933 | 7189 | Normal | 5441 | 22.885 | -70.811 | -73.4372 | -12.5043 | 110.8754 | -74.1281 | 4 | 1 | 25.199 | 1.0633 | -74.1281 | -53.6199 | 1 | 0 | -192 | 0 | 1 | -53.6199
34 | C4 | control | 20 | 1316.5 | 0.28714 | -190.279 | 33 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -178.6245 | 3 | 0 | -190.279 | -190.279 | 5636 | 5636 | 933 | 7189 | Normal | 5636 | 25.156 | -190.279 | -192.245 | -20.2864 | 176.6268 | -191.6934 | 4 | 1 | 25.199 | 1.0633 | -191.6934 | -178.6245 | 1 | 0 | -192 | 0 | 1 | -178.6245
34 | C4 | control | 20 | 1489.5 | 0.32452 | -84.873 | 33 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -67.6196 | 3 | 0 | -84.873 | -84.873 | 5762 | 5762 | 933 | 7189 | Normal | 5762 | 27.232 | -84.873 | -86.4244 | -13.355 | 118.0628 | -87.2947 | 4 | 1 | 25.199 | 1.0633 | -87.2947 | -67.6196 | 1 | 0 | -192 | 0 | 1 | -67.6196
34 | C4 | control | 20 | 1650.7 | 0.26171 | -261.995 | 33 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -253.2359 | 3 | 0 | -261.995 | -261.995 | 5789 | 5789 | 933 | 7189 | Normal | 5789 | 28.494 | -261.995 | -263.4587 | -24.9511 | 216.0384 | -261.8644 | 4 | 1 | 25.199 | 1.0633 | -261.8644 | -253.2359 | 1 | 0 | -192 | 0 | 1 | -253.2359
35 | LCA | size | 30 | 1316.1 | 0.29273 | -174.572 | 35 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -161.1956 | 3 | 0 | -174.572 | -174.572 | 5903 | 5903 | 933 | 7189 | Normal | 5903 | 28.446 | -174.572 | -175.6701 | -19.2007 | 167.4538 | -175.3017 | 4 | 1 | 25.199 | 1.0633 | -175.3017 | -161.1956 | 1 | 0 | -192 | 0 | 1 | -161.1956
35 | LCA | size | 30 | 1489.3 | 0.32498 | -83.645 | 35 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -65.3778 | 3 | 0 | -83.645 | -83.645 | 6176 | 6176 | 933 | 7189 | Normal | 6176 | 30.132 | -83.645 | -83.8953 | -13.1893 | 116.6632 | -85.1863 | 4 | 1 | 25.199 | 1.0633 | -85.1863 | -65.3778 | 1 | 0 | -192 | 0 | 1 | -65.3778
35 | LCA | size | 30 | 1650.7 | 0.27968 | -211.383 | 35 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -198.5276 | 3 | 0 | -211.383 | -211.383 | 6227 | 6227 | 933 | 7189 | Normal | 6227 | 31.703 | -211.383 | -211.4791 | -21.5463 | 187.2715 | -210.412 | 4 | 1 | 25.199 | 1.0633 | -210.412 | -198.5276 | 1 | 0 | -192 | 0 | 1 | -198.5276
35 | LCA | size | 30 | 1800.7 | 0.27891 | -213.545 | 35 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -200.6524 | 3 | 0 | -213.545 | -213.545 | 6263 | 6263 | 933 | 7189 | Normal | 6263 | 32.845 | -213.545 | -213.5331 | -21.6808 | 188.4082 | -212.4103 | 4 | 1 | 25.199 | 1.0633 | -212.4103 | -200.6524 | 1 | 0 | -192 | 0 | 1 | -200.6524
35 | LCA | size | 30 | 1941.2 | 0.28535 | -195.394 | 35 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -181.6113 | 3 | 0 | -195.394 | -195.394 | 6296 | 6296 | 933 | 7189 | Normal | 6296 | 34.559 | -195.394 | -195.2835 | -20.4855 | 178.3084 | -194.5024 | 4 | 1 | 25.199 | 1.0633 | -194.5024 | -181.6113 | 1 | 0 | -192 | 0 | 1 | -181.6113
35 | LCA | size | 30 | 2074.1 | 0.30447 | -141.466 | 35 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -123.0764 | 3 | 0 | -141.466 | -141.466 | 7189 | 7189 | 933 | 7189 | Normal | 7189 | 43.17 | -141.466 | -138.8685 | -16.7902 | 147.0868 | -139.4511 | 4 | 1 | 25.199 | 1.0633 | -139.4511 | -123.0764 | 1 | 0 | -192 | 0 | 1 | -123.0764
35 | LCA | size | 30 | 2217.9 | 0.30226 | -147.72 | 35 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.3104 | 3 | 0 | -147.72 | -147.72 | 6159 | 6159 | 933 | 7189 | Normal | 6159 | 42.351 | -147.72 | -148.022 | -17.3897 | 152.1526 | -148.1356 | 4 | 1 | 25.199 | 1.0633 | -148.1356 | -132.3104 | 1 | 0 | -192 | 0 | 1 | -132.3104
35 | LCA | size | 30 | 2396.2 | 0.28435 | -198.218 | 35 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -190.1355 | 3 | 0 | -198.218 | -198.218 | 4881 | 4881 | 933 | 7189 | Normal | 4881 | 40.709 | -198.218 | -202.8807 | -20.9831 | 182.5129 | -202.5193 | 4 | 1 | 25.199 | 1.0633 | -202.5193 | -190.1355 | 1 | 0 | -192 | 0 | 1 | -190.1355
36 | LCA | size | 25 | 1316.5 | 0.29363 | -172.161 | 35 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -162.7033 | 3 | 0 | -172.161 | -172.161 | 4835 | 4835 | 933 | 7189 | Normal | 4835 | 21.032 | -172.161 | -177.0013 | -19.2879 | 168.1905 | -176.7197 | 4 | 1 | 25.199 | 1.0633 | -176.7197 | -162.7033 | 1 | 0 | -192 | 0 | 1 | -162.7033
36 | LCA | size | 25 | 1489.5 | 0.32546 | -82.445 | 35 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -67.2075 | 3 | 0 | -82.445 | -82.445 | 4945 | 4945 | 933 | 7189 | Normal | 4945 | 22.519 | -82.445 | -86.8635 | -13.3837 | 118.3058 | -86.9071 | 4 | 1 | 25.199 | 1.0633 | -86.9071 | -67.2075 | 1 | 0 | -192 | 0 | 1 | -67.2075
36 | LCA | size | 25 | 1650.5 | 0.28097 | -207.853 | 35 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -199.8541 | 3 | 0 | -207.853 | -207.853 | 4988 | 4988 | 933 | 7189 | Normal | 4988 | 23.691 | -207.853 | -212.1091 | -21.5876 | 187.6202 | -211.6596 | 4 | 1 | 25.199 | 1.0633 | -211.6596 | -199.8541 | 1 | 0 | -192 | 0 | 1 | -199.8541
36 | LCA | size | 25 | 1800.7 | 0.28026 | -209.866 | 35 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -202.013 | 3 | 0 | -209.866 | -209.866 | 4982 | 4982 | 933 | 7189 | Normal | 4982 | 24.648 | -209.866 | -214.1447 | -21.7209 | 188.7467 | -213.6899 | 4 | 1 | 25.199 | 1.0633 | -213.6899 | -202.013 | 1 | 0 | -192 | 0 | 1 | -202.013
36 | LCA | size | 25 | 1941 | 0.28621 | -193.079 | 35 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -183.8787 | 3 | 0 | -193.079 | -193.079 | 5065 | 5065 | 933 | 7189 | Normal | 5065 | 26.038 | -193.079 | -197.0479 | -20.601 | 179.2849 | -196.6349 | 4 | 1 | 25.199 | 1.0633 | -196.6349 | -183.8787 | 1 | 0 | -192 | 0 | 1 | -183.8787
36 | LCA | size | 25 | 2073.9 | 0.30567 | -138.236 | 35 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -123.475 | 3 | 0 | -138.236 | -138.236 | 5796 | 5796 | 933 | 7189 | Normal | 5796 | 32.842 | -138.236 | -139.6771 | -16.8431 | 147.5343 | -139.826 | 4 | 1 | 25.199 | 1.0633 | -139.826 | -123.475 | 1 | 0 | -192 | 0 | 1 | -123.475
36 | LCA | size | 25 | 2217.5 | 0.30378 | -143.557 | 35 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -131.6373 | 3 | 0 | -143.557 | -143.557 | 5028 | 5028 | 933 | 7189 | Normal | 5028 | 32.249 | -143.557 | -147.6634 | -17.3662 | 151.9541 | -147.5026 | 4 | 1 | 25.199 | 1.0633 | -147.5026 | -131.6373 | 1 | 0 | -192 | 0 | 1 | -131.6373
36 | LCA | size | 25 | 2395.3 | 0.28587 | -194.044 | 35 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -189.5139 | 3 | 0 | -194.044 | -194.044 | 4112 | 4112 | 933 | 7189 | Normal | 4112 | 31.212 | -194.044 | -201.9214 | -20.9202 | 181.982 | -201.9347 | 4 | 1 | 25.199 | 1.0633 | -201.9347 | -189.5139 | 1 | 0 | -192 | 0 | 1 | -189.5139
37 | LCA | size | 20 | 1315.4 | 0.29431 | -170.53 | 35 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -163.6874 | 3 | 0 | -170.53 | -170.53 | 4236 | 4236 | 933 | 7189 | Normal | 4236 | 17.67 | -170.53 | -177.8503 | -19.3435 | 168.6604 | -177.6452 | 4 | 1 | 25.199 | 1.0633 | -177.6452 | -163.6874 | 1 | 0 | -192 | 0 | 1 | -163.6874
37 | LCA | size | 20 | 1488.5 | 0.3264 | -80.094 | 35 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -66.819 | 3 | 0 | -80.094 | -80.094 | 4264 | 4264 | 933 | 7189 | Normal | 4264 | 18.759 | -80.094 | -87.2907 | -13.4117 | 118.5423 | -86.5417 | 4 | 1 | 25.199 | 1.0633 | -86.5417 | -66.819 | 1 | 0 | -192 | 0 | 1 | -66.819
37 | LCA | size | 20 | 1649.6 | 0.28184 | -205.673 | 35 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -200.6742 | 3 | 0 | -205.673 | -205.673 | 4351 | 4351 | 933 | 7189 | Normal | 4351 | 19.73 | -205.673 | -212.491 | -21.6126 | 187.8315 | -212.4308 | 4 | 1 | 25.199 | 1.0633 | -212.4308 | -200.6742 | 1 | 0 | -192 | 0 | 1 | -200.6742
37 | LCA | size | 20 | 1799.7 | 0.28155 | -206.517 | 35 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -201.8756 | 3 | 0 | -206.517 | -206.517 | 4295 | 4295 | 933 | 7189 | Normal | 4295 | 20.489 | -206.517 | -213.5779 | -21.6838 | 188.433 | -213.5608 | 4 | 1 | 25.199 | 1.0633 | -213.5608 | -201.8756 | 1 | 0 | -192 | 0 | 1 | -201.8756
37 | LCA | size | 20 | 1940.1 | 0.28778 | -188.949 | 35 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -182.8901 | 3 | 0 | -188.949 | -188.949 | 4335 | 4335 | 933 | 7189 | Normal | 4335 | 21.627 | -188.949 | -195.8361 | -20.5216 | 178.6142 | -195.7051 | 4 | 1 | 25.199 | 1.0633 | -195.7051 | -182.8901 | 1 | 0 | -192 | 0 | 1 | -182.8901
37 | LCA | size | 20 | 2072.9 | 0.30641 | -136.44 | 35 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -124.235 | 3 | 0 | -136.44 | -136.44 | 4992 | 4992 | 933 | 7189 | Normal | 4992 | 27.251 | -136.44 | -140.6811 | -16.9089 | 148.09 | -140.5408 | 4 | 1 | 25.199 | 1.0633 | -140.5408 | -124.235 | 1 | 0 | -192 | 0 | 1 | -124.235
37 | LCA | size | 20 | 2216.2 | 0.30433 | -142.315 | 35 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.828 | 3 | 0 | -142.315 | -142.315 | 4390 | 4390 | 933 | 7189 | Normal | 4390 | 26.728 | -142.315 | -148.9657 | -17.4515 | 152.6749 | -148.6223 | 4 | 1 | 25.199 | 1.0633 | -148.6223 | -132.828 | 1 | 0 | -192 | 0 | 1 | -132.828
37 | LCA | size | 20 | 2393.7 | 0.28644 | -192.721 | 35 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -191.3269 | 3 | 0 | -192.721 | -192.721 | 3562 | 3562 | 933 | 7189 | Normal | 3562 | 25.845 | -192.721 | -203.2907 | -21.0099 | 182.7398 | -203.6399 | 4 | 1 | 25.199 | 1.0633 | -203.6399 | -191.3269 | 1 | 0 | -192 | 0 | 1 | -191.3269
38 | LCA | size | 15 | 1314.6 | 0.29655 | -163.802 | 35 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -161.785 | 3 | 0 | -163.802 | -163.802 | 3274 | 3274 | 933 | 7189 | Normal | 3274 | 14.066 | -163.802 | -175.9526 | -19.2192 | 167.6101 | -175.8561 | 4 | 1 | 25.199 | 1.0633 | -175.8561 | -161.785 | 1 | 0 | -192 | 0 | 1 | -161.785
38 | LCA | size | 15 | 1488.1 | 0.32765 | -76.088 | 35 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -65.5288 | 3 | 0 | -76.088 | -76.088 | 3462 | 3462 | 933 | 7189 | Normal | 3462 | 14.938 | -76.088 | -87.1916 | -13.4052 | 118.4874 | -85.3283 | 4 | 1 | 25.199 | 1.0633 | -85.3283 | -65.5288 | 1 | 0 | -192 | 0 | 1 | -65.5288
38 | LCA | size | 15 | 1649.2 | 0.28354 | -200.482 | 35 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -199.8301 | 3 | 0 | -200.482 | -200.482 | 3546 | 3546 | 933 | 7189 | Normal | 3546 | 15.667 | -200.482 | -211.1361 | -21.5238 | 187.0817 | -211.637 | 4 | 1 | 25.199 | 1.0633 | -211.637 | -199.8301 | 1 | 0 | -192 | 0 | 1 | -199.8301
38 | LCA | size | 15 | 1799.5 | 0.28258 | -203.182 | 35 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -202.5423 | 3 | 0 | -203.182 | -203.182 | 3578 | 3578 | 933 | 7189 | Normal | 3578 | 16.342 | -203.182 | -213.6677 | -21.6897 | 188.4827 | -214.1878 | 4 | 1 | 25.199 | 1.0633 | -214.1878 | -202.5423 | 1 | 0 | -192 | 0 | 1 | -202.5423
38 | LCA | size | 15 | 1939.7 | 0.28893 | -185.275 | 35 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -183.1204 | 3 | 0 | -185.275 | -185.275 | 3586 | 3586 | 933 | 7189 | Normal | 3586 | 17.226 | -185.275 | -195.7188 | -20.514 | 178.5493 | -195.9217 | 4 | 1 | 25.199 | 1.0633 | -195.9217 | -183.1204 | 1 | 0 | -192 | 0 | 1 | -183.1204
38 | LCA | size | 15 | 2072.4 | 0.30723 | -133.678 | 35 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -124.6165 | 3 | 0 | -133.678 | -133.678 | 4152 | 4152 | 933 | 7189 | Normal | 4152 | 21.659 | -133.678 | -141.3738 | -16.9543 | 148.4733 | -140.8995 | 4 | 1 | 25.199 | 1.0633 | -140.8995 | -124.6165 | 1 | 0 | -192 | 0 | 1 | -124.6165
38 | LCA | size | 15 | 2215.8 | 0.30538 | -138.881 | 35 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.6492 | 3 | 0 | -138.881 | -138.881 | 3641 | 3641 | 933 | 7189 | Normal | 3641 | 21.2 | -138.881 | -149.0394 | -17.4564 | 152.7157 | -148.4542 | 4 | 1 | 25.199 | 1.0633 | -148.4542 | -132.6492 | 1 | 0 | -192 | 0 | 1 | -132.6492
38 | LCA | size | 15 | 2392.8 | 0.28846 | -186.601 | 35 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -188.9021 | 3 | 0 | -186.601 | -186.601 | 2956 | 2956 | 933 | 7189 | Normal | 2956 | 20.325 | -186.601 | -200.6674 | -20.8381 | 181.288 | -201.3593 | 4 | 1 | 25.199 | 1.0633 | -201.3593 | -188.9021 | 1 | 0 | -192 | 0 | 1 | -188.9021
39 | LCA | size | 10 | 1314 | 0.29845 | -158.481 | 35 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -161.5261 | 3 | 0 | -158.481 | -158.481 | 2512 | 2512 | 933 | 7189 | Normal | 2512 | 9.855 | -158.481 | -175.5992 | -19.1961 | 167.4146 | -175.6125 | 4 | 1 | 25.199 | 1.0633 | -175.6125 | -161.5261 | 1 | 0 | -192 | 0 | 1 | -161.5261
39 | LCA | size | 10 | 1487.2 | 0.32837 | -74.137 | 35 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -67.8107 | 3 | 0 | -74.137 | -74.137 | 2577 | 2577 | 933 | 7189 | Normal | 2577 | 10.435 | -74.137 | -90.7762 | -13.64 | 120.4712 | -87.4744 | 4 | 1 | 25.199 | 1.0633 | -87.4744 | -67.8107 | 1 | 0 | -192 | 0 | 1 | -67.8107
39 | LCA | size | 10 | 1648.4 | 0.28527 | -195.639 | 35 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -201.7493 | 3 | 0 | -195.639 | -195.639 | 2608 | 2608 | 933 | 7189 | Normal | 2608 | 10.982 | -195.639 | -212.054 | -21.584 | 187.5896 | -213.4419 | 4 | 1 | 25.199 | 1.0633 | -213.4419 | -201.7493 | 1 | 0 | -192 | 0 | 1 | -201.7493
39 | LCA | size | 10 | 1798.4 | 0.28464 | -197.431 | 35 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -203.9386 | 3 | 0 | -197.431 | -197.431 | 2585 | 2585 | 933 | 7189 | Normal | 2585 | 11.391 | -197.431 | -214.0121 | -21.7122 | 188.6733 | -215.501 | 4 | 1 | 25.199 | 1.0633 | -215.501 | -203.9386 | 1 | 0 | -192 | 0 | 1 | -203.9386
39 | LCA | size | 10 | 1938.9 | 0.29083 | -179.973 | 35 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -184.3783 | 3 | 0 | -179.973 | -179.973 | 2617 | 2617 | 933 | 7189 | Normal | 2617 | 12.035 | -179.973 | -196.3234 | -20.5536 | 178.8839 | -197.1048 | 4 | 1 | 25.199 | 1.0633 | -197.1048 | -184.3783 | 1 | 0 | -192 | 0 | 1 | -184.3783
39 | LCA | size | 10 | 2071.4 | 0.30933 | -127.822 | 35 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -124.1505 | 3 | 0 | -127.822 | -127.822 | 3017 | 3017 | 933 | 7189 | Normal | 3017 | 15.148 | -127.822 | -141.5054 | -16.9629 | 148.5461 | -140.4613 | 4 | 1 | 25.199 | 1.0633 | -140.4613 | -124.1505 | 1 | 0 | -192 | 0 | 1 | -124.1505
39 | LCA | size | 10 | 2214.4 | 0.30743 | -133.176 | 35 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.1221 | 3 | 0 | -133.176 | -133.176 | 2701 | 2701 | 933 | 7189 | Normal | 2701 | 14.833 | -133.176 | -148.934 | -17.4495 | 152.6573 | -147.9585 | 4 | 1 | 25.199 | 1.0633 | -147.9585 | -132.1221 | 1 | 0 | -192 | 0 | 1 | -132.1221
39 | LCA | size | 10 | 2391.2 | 0.29136 | -178.481 | 35 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -186.2643 | 3 | 0 | -178.481 | -178.481 | 2240 | 2240 | 933 | 7189 | Normal | 2240 | 14.221 | -178.481 | -197.7481 | -20.6469 | 179.6724 | -198.8785 | 4 | 1 | 25.199 | 1.0633 | -198.8785 | -186.2643 | 1 | 0 | -192 | 0 | 1 | -186.2643
40 | LCA | size | 5 | 1314 | 0.30278 | -146.179 | 35 | C21 Alkane | C21 | Alkane | LCA | 1 |  |  |  | -165.6616 | 3 | 0 | -146.179 | -146.179 | 1098 | 1098 | 933 | 7189 | Normal | 1098 | 3.954 | -146.179 | -178.815 | -19.4067 | 169.1943 | -179.502 | 4 | 1 | 25.199 | 1.0633 | -179.502 | -165.6616 | 1 | 0 | -192 | 0 | 1 | -165.6616
40 | LCA | size | 5 | 1487 | 0.33264 | -61.991 | 35 | C23 Alkane | C23 | Alkane | LCA | 1 |  |  |  | -67.3869 | 3 | 0 | -61.991 | -61.991 | 1116 | 1116 | 933 | 7189 | Normal | 1116 | 4.231 | -61.991 | -94.3221 | -13.8723 | 122.4336 | -87.0759 | 4 | 1 | 25.199 | 1.0633 | -87.0759 | -67.3869 | 1 | 0 | -192 | 0 | 1 | -67.3869
40 | LCA | size | 5 | 1648 | 0.29211 | -176.292 | 35 | C25 Alkane | C25 | Alkane | LCA | 1 |  |  |  | -199.9885 | 3 | 0 | -176.292 | -176.292 | 1131 | 1131 | 933 | 7189 | Normal | 1131 | 4.487 | -176.292 | -208.3727 | -21.3428 | 185.5523 | -211.786 | 4 | 1 | 25.199 | 1.0633 | -211.786 | -199.9885 | 1 | 0 | -192 | 0 | 1 | -199.9885
40 | LCA | size | 5 | 1798 | 0.29097 | -179.502 | 35 | C27 Alkane | C27 | Alkane | LCA | 1 |  |  |  | -203.7406 | 3 | 0 | -179.502 | -179.502 | 1130 | 1130 | 933 | 7189 | Normal | 1130 | 4.7 | -179.502 | -211.5993 | -21.5542 | 187.338 | -215.3148 | 4 | 1 | 25.199 | 1.0633 | -215.3148 | -203.7406 | 1 | 0 | -192 | 0 | 1 | -203.7406
40 | LCA | size | 5 | 1938.3 | 0.29678 | -163.123 | 35 | C29 Alkane | C29 | Alkane | LCA | 1 |  |  |  | -184.3538 | 3 | 0 | -163.123 | -163.123 | 1147 | 1147 | 933 | 7189 | Normal | 1147 | 5.018 | -163.123 | -194.9403 | -20.463 | 178.1185 | -197.0818 | 4 | 1 | 25.199 | 1.0633 | -197.0818 | -184.3538 | 1 | 0 | -192 | 0 | 1 | -184.3538
40 | LCA | size | 5 | 2070.6 | 0.31325 | -116.654 | 35 | C31 Alkane | C31 | Alkane | LCA | 1 |  |  |  | -127.0295 | 3 | 0 | -116.654 | -116.654 | 1356 | 1356 | 933 | 7189 | Normal | 1356 | 6.339 | -116.654 | -145.3327 | -17.2136 | 150.6643 | -143.1689 | 4 | 1 | 25.199 | 1.0633 | -143.1689 | -127.0295 | 1 | 0 | -192 | 0 | 1 | -127.0295
40 | LCA | size | 5 | 2213.1 | 0.31224 | -119.504 | 35 | C33 Alkane | C33 | Alkane | LCA | 1 |  |  |  | -132.1147 | 3 | 0 | -119.504 | -119.504 | 1240 | 1240 | 933 | 7189 | Normal | 1240 | 6.259 | -119.504 | -149.8595 | -17.5101 | 153.1695 | -147.9515 | 4 | 1 | 25.199 | 1.0633 | -147.9515 | -132.1147 | 1 | 0 | -192 | 0 | 1 | -132.1147
40 | LCA | size | 5 | 2389.3 | 0.29674 | -163.222 | 35 | C35 Alkane | C35 | Alkane | LCA | 1 |  |  |  | -186.2487 | 3 | 0 | -163.222 | -163.222 | 1064 | 1064 | 933 | 7189 | Normal | 1064 | 6.03 | -163.222 | -196.4478 | -20.5617 | 178.9527 | -198.8639 | 4 | 1 | 25.199 | 1.0633 | -198.8639 | -186.2487 | 1 | 0 | -192 | 0 | 1 | -186.2487
41 | ACAL | standard_control_drift | 20 | 1650.3 | 0.26455 | -254.273 | 41 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -254.1595 | 3 | 0 | -254.273 | -254.273 | 4101 | 4101 | 933 | 7189 | Normal | 4101 | 18.468 | -254.273 | -262.2006 | -24.8687 | 215.3421 | -262.7331 | 5 | 1 | 25.199 | 1.0633 | -262.7331 | -254.1595 | 1 | 0 | -192 | 0 | 1 | -254.1595
41 | ACAL | standard_control_drift | 20 | 1800.5 | 0.29111 | -179.398 | 41 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -174.0152 | 3 | 0 | -179.398 | -179.398 | 4073 | 4073 | 933 | 7189 | Normal | 4073 | 19.106 | -179.398 | -187.454 | -19.9726 | 173.9754 | -187.3583 | 5 | 1 | 25.199 | 1.0633 | -187.3583 | -174.0152 | 1 | 0 | -192 | 0 | 1 | -174.0152
41 | ACAL | standard_control_drift | 20 | 1940.8 | 0.31974 | -98.692 | 41 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -87.1261 | 3 | 0 | -98.692 | -98.692 | 4153 | 4153 | 933 | 7189 | Normal | 4153 | 20.618 | -98.692 | -106.3833 | -14.6623 | 129.1086 | -105.6403 | 5 | 1 | 25.199 | 1.0633 | -105.6403 | -87.1261 | 1 | 0 | -192 | 0 | 1 | -87.1261
42 | ACAL | standard_control_drift | 20 | 1649 | 0.26474 | -253.575 | 41 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -254.4837 | 3 | 0 | -253.575 | -253.575 | 3939 | 3939 | 933 | 7189 | Normal | 3939 | 18.121 | -253.575 | -262.2583 | -24.8724 | 215.3741 | -263.038 | 5 | 1 | 25.199 | 1.0633 | -263.038 | -254.4837 | 1 | 0 | -192 | 0 | 1 | -254.4837
42 | ACAL | standard_control_drift | 20 | 1799.3 | 0.29134 | -178.587 | 41 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -173.823 | 3 | 0 | -178.587 | -178.587 | 3946 | 3946 | 933 | 7189 | Normal | 3946 | 18.745 | -178.587 | -187.237 | -19.9584 | 173.8552 | -187.1776 | 5 | 1 | 25.199 | 1.0633 | -187.1776 | -173.823 | 1 | 0 | -192 | 0 | 1 | -173.823
42 | ACAL | standard_control_drift | 20 | 1939.5 | 0.3199 | -98.045 | 41 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.7117 | 3 | 0 | -98.045 | -98.045 | 4080 | 4080 | 933 | 7189 | Normal | 4080 | 20.227 | -98.045 | -106.0688 | -14.6417 | 128.9346 | -105.2506 | 5 | 1 | 25.199 | 1.0633 | -105.2506 | -86.7117 | 1 | 0 | -192 | 0 | 1 | -86.7117
43 | C4 | control | 10 | 938.6 | 0.31172 | -121.038 | 43 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -119.2589 | 3 | 0 | -121.038 | -121.038 | 2628 | 2628 | 933 | 7189 | Normal | 2628 | 8.97 | -121.038 | -137.3097 | -16.6881 | 146.2241 | -135.8608 | 5 | 1 | 25.199 | 1.0633 | -135.8608 | -119.2589 | 1 | 0 | -192 | 0 | 1 | -119.2589
43 | C4 | control | 10 | 1129.9 | 0.3328 | -61.616 | 43 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -52.9451 | 3 | 0 | -61.616 | -61.616 | 2770 | 2770 | 933 | 7189 | Normal | 2770 | 10.165 | -61.616 | -76.901 | -12.7312 | 112.7923 | -73.4935 | 5 | 1 | 25.199 | 1.0633 | -73.4935 | -52.9451 | 1 | 0 | -192 | 0 | 1 | -52.9451
43 | C4 | control | 10 | 1314.6 | 0.29142 | -178.275 | 43 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -179.8742 | 3 | 0 | -178.275 | -178.275 | 2944 | 2944 | 933 | 7189 | Normal | 2944 | 11.379 | -178.275 | -192.4177 | -20.2977 | 176.7224 | -192.8687 | 5 | 1 | 25.199 | 1.0633 | -192.8687 | -179.8742 | 1 | 0 | -192 | 0 | 1 | -179.8742
43 | C4 | control | 10 | 1487.7 | 0.32774 | -75.882 | 43 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -67.1605 | 3 | 0 | -75.882 | -75.882 | 3054 | 3054 | 933 | 7189 | Normal | 3054 | 12.491 | -75.882 | -89.3369 | -13.5457 | 119.6747 | -86.8629 | 5 | 1 | 25.199 | 1.0633 | -86.8629 | -67.1605 | 1 | 0 | -192 | 0 | 1 | -67.1605
43 | C4 | control | 10 | 1648.8 | 0.26661 | -248.251 | 43 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -255.2992 | 3 | 0 | -248.251 | -248.251 | 3081 | 3081 | 933 | 7189 | Normal | 3081 | 13.149 | -248.251 | -261.5408 | -24.8254 | 214.977 | -263.805 | 5 | 1 | 25.199 | 1.0633 | -263.805 | -255.2992 | 1 | 0 | -192 | 0 | 1 | -255.2992
44 | C4 | control | 10 | 937.2 | 0.31255 | -118.802 | 43 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -117.7552 | 3 | 0 | -118.802 | -118.802 | 2494 | 2494 | 933 | 7189 | Normal | 2494 | 9.096 | -118.802 | -136.0551 | -16.6059 | 145.5298 | -134.4466 | 5 | 1 | 25.199 | 1.0633 | -134.4466 | -117.7552 | 1 | 0 | -192 | 0 | 1 | -117.7552
44 | C4 | control | 10 | 1128.8 | 0.33352 | -59.692 | 43 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -51.7321 | 3 | 0 | -59.692 | -59.692 | 2599 | 2599 | 933 | 7189 | Normal | 2599 | 10.38 | -59.692 | -76.1718 | -12.6834 | 112.3888 | -72.3527 | 5 | 1 | 25.199 | 1.0633 | -72.3527 | -51.7321 | 1 | 0 | -192 | 0 | 1 | -51.7321
44 | C4 | control | 10 | 1314 | 0.29194 | -176.924 | 43 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -179.7932 | 3 | 0 | -176.924 | -176.924 | 2764 | 2764 | 933 | 7189 | Normal | 2764 | 11.638 | -176.924 | -192.2497 | -20.2867 | 176.6294 | -192.7925 | 5 | 1 | 25.199 | 1.0633 | -192.7925 | -179.7932 | 1 | 0 | -192 | 0 | 1 | -179.7932
44 | C4 | control | 10 | 1487.2 | 0.32782 | -75.759 | 43 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -67.3986 | 3 | 0 | -75.759 | -75.759 | 2979 | 2979 | 933 | 7189 | Normal | 2979 | 12.716 | -75.759 | -89.6801 | -13.5682 | 119.8646 | -87.0868 | 5 | 1 | 25.199 | 1.0633 | -87.0868 | -67.3986 | 1 | 0 | -192 | 0 | 1 | -67.3986
44 | C4 | control | 10 | 1648.6 | 0.26698 | -247.284 | 43 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -254.4444 | 3 | 0 | -247.284 | -247.284 | 3058 | 3058 | 933 | 7189 | Normal | 3058 | 13.425 | -247.284 | -260.7143 | -24.7713 | 214.5196 | -263.0011 | 5 | 1 | 25.199 | 1.0633 | -263.0011 | -254.4444 | 1 | 0 | -192 | 0 | 1 | -254.4444
45 | C4 | control | 5 | 938.6 | 0.3154 | -110.165 | 43 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -120.8211 | 3 | 0 | -110.165 | -110.165 | 1272 | 1272 | 933 | 7189 | Normal | 1272 | 4.19 | -110.165 | -140.0428 | -16.8671 | 147.7367 | -137.33 | 5 | 1 | 25.199 | 1.0633 | -137.33 | -120.8211 | 1 | 0 | -192 | 0 | 1 | -120.8211
45 | C4 | control | 5 | 1129.6 | 0.33498 | -54.919 | 43 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -55.8195 | 3 | 0 | -54.919 | -54.919 | 1378 | 1378 | 933 | 7189 | Normal | 1378 | 4.775 | -54.919 | -83.2959 | -13.15 | 116.3314 | -76.1969 | 5 | 1 | 25.199 | 1.0633 | -76.1969 | -55.8195 | 1 | 0 | -192 | 0 | 1 | -55.8195
45 | C4 | control | 5 | 1314.4 | 0.29599 | -164.939 | 43 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -181.0062 | 3 | 0 | -164.939 | -164.939 | 1448 | 1448 | 933 | 7189 | Normal | 1448 | 5.289 | -164.939 | -192.3868 | -20.2957 | 176.7053 | -193.9333 | 5 | 1 | 25.199 | 1.0633 | -193.9333 | -181.0062 | 1 | 0 | -192 | 0 | 1 | -181.0062
45 | C4 | control | 5 | 1487.2 | 0.33125 | -65.452 | 43 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -66.5832 | 3 | 0 | -65.452 | -65.452 | 1496 | 1496 | 933 | 7189 | Normal | 1496 | 5.698 | -65.452 | -92.2883 | -13.7391 | 121.3081 | -86.32 | 5 | 1 | 25.199 | 1.0633 | -86.32 | -66.5832 | 1 | 0 | -192 | 0 | 1 | -66.5832
45 | C4 | control | 5 | 1648.4 | 0.27303 | -229.698 | 43 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -254.4674 | 3 | 0 | -229.698 | -229.698 | 1484 | 1484 | 933 | 7189 | Normal | 1484 | 5.906 | -229.698 | -256.6854 | -24.5074 | 212.2898 | -263.0226 | 5 | 1 | 25.199 | 1.0633 | -263.0226 | -254.4674 | 1 | 0 | -192 | 0 | 1 | -254.4674
46 | C4 | control | 5 | 936.9 | 0.31564 | -109.657 | 43 | C17 Alkane | C17 | Alkane | C4 | 1 | -117.8 | 2.1 | 7 | -118.887 | 3 | 0 | -109.657 | -109.657 | 1363 | 1363 | 933 | 7189 | Normal | 1363 | 4.558 | -109.657 | -138.2392 | -16.7489 | 146.7385 | -135.511 | 5 | 1 | 25.199 | 1.0633 | -135.511 | -118.887 | 1 | 0 | -192 | 0 | 1 | -118.887
46 | C4 | control | 5 | 1128.6 | 0.33531 | -54.183 | 43 | C19 Alkane | C19 | Alkane | C4 | 1 | -56.3 | 1 | 5 | -54.4011 | 3 | 0 | -54.183 | -54.183 | 1430 | 1430 | 933 | 7189 | Normal | 1430 | 5.023 | -54.183 | -81.8654 | -13.0563 | 115.5397 | -74.8629 | 5 | 1 | 25.199 | 1.0633 | -74.8629 | -54.4011 | 1 | 0 | -192 | 0 | 1 | -54.4011
46 | C4 | control | 5 | 1313.6 | 0.29537 | -166.831 | 43 | C21 Alkane | C21 | Alkane | C4 | 1 | -177.8 | 1.5 | 6 | -182.0627 | 3 | 0 | -166.831 | -166.831 | 1519 | 1519 | 933 | 7189 | Normal | 1519 | 5.531 | -166.831 | -193.3813 | -20.3609 | 177.2556 | -194.927 | 5 | 1 | 25.199 | 1.0633 | -194.927 | -182.0627 | 1 | 0 | -192 | 0 | 1 | -182.0627
46 | C4 | control | 5 | 1486.6 | 0.32983 | -69.627 | 43 | C23 Alkane | C23 | Alkane | C4 | 1 | -67.2 | 1.1 | 6 | -70.5447 | 3 | 0 | -69.627 | -69.627 | 1572 | 1572 | 933 | 7189 | Normal | 1572 | 6.007 | -69.627 | -95.5342 | -13.9517 | 123.1044 | -90.0457 | 5 | 1 | 25.199 | 1.0633 | -90.0457 | -70.5447 | 1 | 0 | -192 | 0 | 1 | -70.5447
46 | C4 | control | 5 | 1647.8 | 0.27284 | -230.386 | 43 | C25 Alkane | C25 | Alkane | C4 | 1 | -254.1 | 1.5 | 5 | -253.6506 | 3 | 0 | -230.386 | -230.386 | 1573 | 1573 | 933 | 7189 | Normal | 1573 | 6.33 | -230.386 | -256.2813 | -24.4809 | 212.0662 | -262.2544 | 5 | 1 | 25.199 | 1.0633 | -262.2544 | -253.6506 | 1 | 0 | -192 | 0 | 1 | -253.6506
47 | BOF1-41 | sample |  | 1644.6 | 0.27792 | -216.186 | 47 | C25 Alkane | C25 | Alkane | [5 empty] | -232.1179 | 3 | 0 | -216.186 | -216.186 | 1930 | 1930 | 933 | 7189 | Normal | 1930 | 8.853 | -216.186 | -238.2461 | -23.2996 | 202.085 | -242.0032 | 5 | 1 | 25.199 | 1.0633 | -242.0032 | -232.1179 | 1 | 0 | -192 | 0 | 1 | -232.1179
47 | BOF1-41 | sample |  | 1795.7 | 0.27921 | -212.574 | 47 | C27 Alkane | C27 | Alkane | [5 empty] | -227.3346 | 3 | 0 | -212.574 | -212.574 | 1985 | 1985 | 933 | 7189 | Normal | 1985 | 9.55 | -212.574 | -234.1072 | -23.0285 | 199.7945 | -237.5046 | 5 | 1 | 25.199 | 1.0633 | -237.5046 | -227.3346 | 1 | 0 | -192 | 0 | 1 | -227.3346
47 | BOF1-41 | sample |  | 1936.4 | 0.27687 | -219.17 | 47 | C29 Alkane | C29 | Alkane | [5 empty] | -236.9837 | 3 | 0 | -219.17 | -219.17 | 1822 | 1822 | 933 | 7189 | Normal | 1822 | 9.016 | -219.17 | -242.3099 | -23.5658 | 204.334 | -246.5795 | 5 | 1 | 25.199 | 1.0633 | -246.5795 | -236.9837 | 1 | 0 | -192 | 0 | 1 | -236.9837
48 | BOF1-41 | sample |  | 1482.4 | 0.27967 | -211.056 | 47 | C23 Alkane | C23 | Alkane | [5 empty] | -235.3998 | 3 | 0 | -211.056 | -211.056 | 1363 | 1363 | 933 | 7189 | Normal | 1363 | 5.278 | -211.056 | -239.6382 | -23.3908 | 202.8554 | -245.0899 | 5 | 1 | 25.199 | 1.0633 | -245.0899 | -235.3998 | 1 | 0 | -192 | 0 | 1 | -235.3998
48 | BOF1-41 | sample |  | 1646.1 | 0.27539 | -223.113 | 47 | C25 Alkane | C25 | Alkane | [5 empty] | -233.4319 | 3 | 0 | -223.113 | -223.113 | 2474 | 2474 | 933 | 7189 | Normal | 2474 | 10.462 | -223.113 | -240.517 | -23.4483 | 203.3418 | -243.2391 | 5 | 1 | 25.199 | 1.0633 | -243.2391 | -233.4319 | 1 | 0 | -192 | 0 | 1 | -233.4319
48 | BOF1-41 | sample |  | 1797.4 | 0.27712 | -218.251 | 47 | C27 Alkane | C27 | Alkane | [5 empty] | -227.2394 | 3 | 0 | -218.251 | -218.251 | 2554 | 2554 | 933 | 7189 | Normal | 2554 | 11.376 | -218.251 | -235.0583 | -23.0908 | 200.3208 | -237.415 | 5 | 1 | 25.199 | 1.0633 | -237.415 | -227.2394 | 1 | 0 | -192 | 0 | 1 | -227.2394
48 | BOF1-41 | sample |  | 1938.1 | 0.27467 | -225.166 | 47 | C29 Alkane | C29 | Alkane | [5 empty] | -236.5478 | 3 | 0 | -225.166 | -225.166 | 2395 | 2395 | 933 | 7189 | Normal | 2395 | 11.045 | -225.166 | -243.1785 | -23.6227 | 204.8148 | -246.1695 | 5 | 1 | 25.199 | 1.0633 | -246.1695 | -236.5478 | 1 | 0 | -192 | 0 | 1 | -236.5478
49 | NE 0-5 cm | sample |  | 1796.6 | 0.27527 | -223.958 | 49 | C27 Alkane | C27 | Alkane | [5 empty] | -239.9847 | 3 | 0 | -223.958 | -223.958 | 1995 | 1995 | 933 | 7189 | Normal | 1995 | 8.538 | -223.958 | -245.397 | -23.768 | 206.0426 | -249.4019 | 5 | 1 | 25.199 | 1.0633 | -249.4019 | -239.9847 | 1 | 0 | -192 | 0 | 1 | -239.9847
49 | NE 0-5 cm | sample |  | 1939.3 | 0.2688 | -242.186 | 49 | C29 Alkane | C29 | Alkane | [5 empty] | -236.6008 | 3 | 0 | -242.186 | -242.186 | 4910 | 4910 | 933 | 7189 | Normal | 4910 | 25.073 | -242.186 | -246.7377 | -23.8558 | 206.7845 | -246.2193 | 5 | 1 | 25.199 | 1.0633 | -246.2193 | -236.6008 | 1 | 0 | -192 | 0 | 1 | -236.6008
49 | NE 0-5 cm | sample |  | 2072.2 | 0.26601 | -250.052 | 49 | C31 Alkane | C31 | Alkane | [5 empty] | -242.4617 | 3 | 0 | -250.052 | -250.052 | 5406 | 5406 | 933 | 7189 | Normal | 5406 | 30.143 | -250.052 | -252.7992 | -24.2528 | 210.1391 | -251.7315 | 5 | 1 | 25.199 | 1.0633 | -251.7315 | -242.4617 | 1 | 0 | -192 | 0 | 1 | -242.4617
49 | NE 0-5 cm | sample |  | 2215.2 | 0.26988 | -239.154 | 49 | C33 Alkane | C33 | Alkane | [5 empty] | -239.3933 | 3 | 0 | -239.154 | -239.154 | 3876 | 3876 | 933 | 7189 | Normal | 3876 | 23.059 | -239.154 | -248.1396 | -23.9476 | 207.5604 | -248.8456 | 5 | 1 | 25.199 | 1.0633 | -248.8456 | -239.3933 | 1 | 0 | -192 | 0 | 1 | -239.3933
50 | NE 0-5 cm | sample |  | 1797.4 | 0.27452 | -226.074 | 49 | C27 Alkane | C27 | Alkane | [5 empty] | -242.2406 | 3 | 0 | -226.074 | -226.074 | 2004 | 2004 | 933 | 7189 | Normal | 2004 | 8.498 | -226.074 | -247.4286 | -23.9011 | 207.1669 | -251.5235 | 5 | 1 | 25.199 | 1.0633 | -251.5235 | -242.2406 | 1 | 0 | -192 | 0 | 1 | -242.2406
50 | NE 0-5 cm | sample |  | 1940.1 | 0.26841 | -243.314 | 49 | C29 Alkane | C29 | Alkane | [5 empty] | -237.8531 | 3 | 0 | -243.314 | -243.314 | 4899 | 4899 | 933 | 7189 | Normal | 4899 | 25.075 | -243.314 | -247.9077 | -23.9324 | 207.4321 | -247.3972 | 5 | 1 | 25.199 | 1.0633 | -247.3972 | -237.8531 | 1 | 0 | -192 | 0 | 1 | -237.8531
50 | NE 0-5 cm | sample |  | 2073.1 | 0.26571 | -250.906 | 49 | C31 Alkane | C31 | Alkane | [5 empty] | -243.2853 | 3 | 0 | -250.906 | -250.906 | 5422 | 5422 | 933 | 7189 | Normal | 5422 | 30.148 | -250.906 | -253.5978 | -24.3052 | 210.5811 | -252.506 | 5 | 1 | 25.199 | 1.0633 | -252.506 | -243.2853 | 1 | 0 | -192 | 0 | 1 | -243.2853
50 | NE 0-5 cm | sample |  | 2216.2 | 0.26984 | -239.276 | 49 | C33 Alkane | C33 | Alkane | [5 empty] | -239.3791 | 3 | 0 | -239.276 | -239.276 | 3898 | 3898 | 933 | 7189 | Normal | 3898 | 23.089 | -239.276 | -248.1555 | -23.9487 | 207.5692 | -248.8323 | 5 | 1 | 25.199 | 1.0633 | -248.8323 | -239.3791 | 1 | 0 | -192 | 0 | 1 | -239.3791
51 | ACAL | standard_control_drift | 20 | 1650.1 | 0.26485 | -252.958 | 51 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -253.9696 | 3 | 0 | -252.958 | -252.958 | 3917 | 3917 | 933 | 7189 | Normal | 3917 | 18.264 | -252.958 | -261.7463 | -24.8389 | 215.0907 | -262.5544 | 5 | 1 | 25.199 | 1.0633 | -262.5544 | -253.9696 | 1 | 0 | -192 | 0 | 1 | -253.9696
51 | ACAL | standard_control_drift | 20 | 1800.1 | 0.29133 | -178.275 | 51 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -173.4388 | 3 | 0 | -178.275 | -178.275 | 3955 | 3955 | 933 | 7189 | Normal | 3955 | 18.792 | -178.275 | -186.8823 | -19.9352 | 173.6589 | -186.8163 | 5 | 1 | 25.199 | 1.0633 | -186.8163 | -173.4388 | 1 | 0 | -192 | 0 | 1 | -173.4388
51 | ACAL | standard_control_drift | 20 | 1940.6 | 0.31999 | -97.428 | 51 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -86.073 | 3 | 0 | -97.428 | -97.428 | 4074 | 4074 | 933 | 7189 | Normal | 4074 | 20.249 | -97.428 | -105.4794 | -14.6031 | 128.6084 | -104.6498 | 5 | 1 | 25.199 | 1.0633 | -104.6498 | -86.073 | 1 | 0 | -192 | 0 | 1 | -86.073
52 | ACAL | standard_control_drift | 20 | 1650.1 | 0.2644 | -254.37 | 51 | C25 Alkane | C25 | Alkane | ACAL | 1 | -263.6 |  |  | -254.2376 | 3 | 0 | -254.37 | -254.37 | 4105 | 4105 | 933 | 7189 | Normal | 4105 | 18.313 | -254.37 | -262.2793 | -24.8738 | 215.3857 | -262.8065 | 5 | 1 | 25.199 | 1.0633 | -262.8065 | -254.2376 | 1 | 0 | -192 | 0 | 1 | -254.2376
52 | ACAL | standard_control_drift | 20 | 1800.1 | 0.29135 | -178.361 | 51 | C27 Alkane | C27 | Alkane | ACAL | 1 | -172.8 |  |  | -173.0025 | 3 | 0 | -178.361 | -178.361 | 4054 | 4054 | 933 | 7189 | Normal | 4054 | 18.929 | -178.361 | -186.5047 | -19.9104 | 173.45 | -186.4059 | 5 | 1 | 25.199 | 1.0633 | -186.4059 | -173.0025 | 1 | 0 | -192 | 0 | 1 | -173.0025
52 | ACAL | standard_control_drift | 20 | 1940.6 | 0.31929 | -99.589 | 51 | C29 Alkane | C29 | Alkane | ACAL | 1 | -85.4 |  |  | -88.2288 | 3 | 0 | -99.589 | -99.589 | 4116 | 4116 | 933 | 7189 | Normal | 4116 | 20.448 | -99.589 | -107.4481 | -14.7321 | 129.6979 | -106.6774 | 5 | 1 | 25.199 | 1.0633 | -106.6774 | -88.2288 | 1 | 0 | -192 | 0 | 1 | -88.2288

## Standard Info
std_mix | comp | class | relative_concentration | dD_known | dD_known_sd | dD_known_n
CAL | C20 | FAME | 1 | -4.9 | 1 | 118
CAL | C30 | FAME | 1 | -189.4 | 2 | 5
CAL | C24 | FAME | 1 | -179.3 | 1.7 | 5
CAL | C31 | Alkane | 1 | -271.9 | 2 | 9
C4 | C17 | Alkane | 1 | -117.8 | 2.1 | 7
C4 | C19 | Alkane | 1 | -56.3 | 1 | 5
C4 | C21 | Alkane | 1 | -177.8 | 1.5 | 6
C4 | C23 | Alkane | 1 | -67.2 | 1.1 | 6
C4 | C25 | Alkane | 1 | -254.1 | 1.5 | 5
LCF | C24 | FAME | 1 | -204.2
LCF | C26 | FAME | 1 | -146.6
LCF | C28 | FAME | 1 | -161.1
LCF | C30 | FAME | 1 | -197.9
LCF | C32 | FAME | 1 | -154.5
F8 | C14 | FAME | 1 | -231.2 | 1.4 | 4
F8 | C14 | FAEE | 1 | -231.2 | 2.7 | 7
F8 | C16 | FAME | 1 | -166.8 | 1.7 | 4
F8 | C16 | FAEE | 1 | -211 | 1.7 | 4
F8 | C18 | FAME | 1 | -206.2 | 1.7 | 5
F8 | C18 | FAEE | 1 | -214.2 | 0.7 | 4
F8 | C20 | FAME | 1 | -166.7 | 0.3 | 3
F8 | C20 | FAEE | 1 | -195.5 | 1.2 | 4
A7 | C16 | Alkane | 1 | -9.1 | 1.4 | 7
A7 | C17 | Alkane | 1 | -117.5 | 2.1 | 8
A7 | C18 | Alkane | 1 | -52 | 1.1 | 5
A7 | C19 | Alkane | 1 | -56.3 | 1 | 5
A7 | C20 | Alkane | 1 | -177.6 | 1.1 | 5
A7 | C21 | Alkane | 1 | -177.8 | 1.5 | 6
A7 | C22 | Alkane | 1 | -75.8 | 1.8 | 5
A7 | C23 | Alkane | 1 | -67.2 | 1 | 6
A7 | C24 | Alkane | 1 | -29.7 | 1.5 | 6
A7 | C25 | Alkane | 1 | -263.6 | 2.2 | 5
A7 | C26 | Alkane | 1 | -45.9 | 1 | 5
A7 | C27 | Alkane | 1 | -172.8 | 1.6 | 6
A7 | C28 | Alkane | 1 | -36.8 | 1.3 | 5
A7 | C29 | Alkane | 1 | -177.8 | 1.3 | 9
A7 | C30 | Alkane | 1 | -213.6 | 2.4 | 6
PA2 | PAME | BPCA | 1 | -81.9 | 1.2 | 4
Jeff QAQC | C25 | Alkane | 1 | -254.1 | 1.5 | 5
Jeff QAQC | C28 | Alkane | 1 | -36.8 | 1.3 | 5
Jeff QAQC | C29 | Alkane | 1 | -177.8 | 1.3 | 5
C20 + C28 FAMEs | C20 | FAME | 1 | -166.7 | 0.3 | 3
C20 + C28 FAMEs | C28 | FAME | 1
C20 + C30 FAMEs | C20 | FAME | 1 | -4.9 | 1 | 118
C20 + C30 FAMEs | C30 | FAME | 1 | -189.4 | 2 | 5
C4 Mix | C17 | Alkane | 1 | -117.8 | 2.1 | 7
C4 Mix | C19 | Alkane | 1 | -56.3 | 1 | 5
C4 Mix | C21 | Alkane | 1 | -177.8 | 1.5 | 6
C4 Mix | C23 | Alkane | 1 | -67.2 | 1.1 | 6
C4 Mix | C25 | Alkane | 1 | -254.1 | 1.5 | 5
C18 + C24 FAMEs | C18 | FAME | 1 | -206.2 | 1.7 | 5
C18 + C24 FAMEs | C24 | FAME | 1 | -179.3 | 1.7 | 5
C27 + C31 Alkanes | C27 | Alkane | 1 | -172.8 | 1.6 | 6
C27 + C31 Alkanes | C31 | Alkane | 1 | -271.9 | 2 | 9
C22 + C25 + C35 Alkanes | C22 | Alkane | 1 | -81.3 | 1.8 | 5
C22 + C25 + C35 Alkanes | C25 | Alkane | 1 | -263 | 2.5 | 6
C22 + C25 + C35 Alkanes | C35 | Alkane | 1 | -179.3 | 1.9 | 4
ACAL | C23 | Alkane | 1 | -67.2
ACAL | C25 | Alkane | 1 | -263.6
ACAL | C27 | Alkane | 1 | -172.8
ACAL | C29 | Alkane | 1 | -85.4
ACAL | C31 | Alkane | 1 | -271.9
LCA | C21 | Alkane | 1
LCA | C23 | Alkane | 1
LCA | C25 | Alkane | 1
LCA | C27 | Alkane | 1
LCA | C29 | Alkane | 1
LCA | C31 | Alkane | 1
LCA | C33 | Alkane | 1
LCA | C35 | Alkane | 1

## Initials
variable | initial_value
compound_option | Assigned by IRMS export in Component/comp column.
first_correction | Drift
second_correction | Size
third_correction | Scale Normalization
drift_option | No drift correction (use this when there is no apparent drift or if you use bracketed scale normalization)
drift_comp | Mean drift of all compounds
size_option | Peak Height (amplitude, mV)
size_model_type | Log-transformed, composition-scaled linear model
size_cutoff
size_toosmall_peak_option | Remove these from results.
size_small_peak_option | Remove observed size effect.
size_normal_peak_option | Remove observed size effect.
size_large_peak_option | Remove size effect with 'Normal' size effect function.
acceptable_peak_units | Peak Height (amplitude, mV)
largest_acceptable_peak
smallest_acceptable_peak
normalization_option | Linear interpolation between adjacent normalization standards (use this if drift-correction is untenable)
id1 | ACAL
normalization_comps | C23 Alkane
normalization_comps | C31 Alkane
derivatization_option | Template-defined derivative δD.

## Headers
ingest_header | export_header
id1 | Identifier.1
id2 | Identifier.2
conc | Preparation
dD_raw | d.2H.1H
peak | Peak.Nr.
area | Area.All
ampl | Ampl..2
raw_R | rR.3H2.2H2
rt | Rt
row | Row
comp | Component

## Retention Times
comp_rt_table | rt_table | window
C17 Alkane | 768.2 | 10
C19 Alkane | 895.7 | 10
C20 FAME | 1217.5 | 10
C21 Alkane | 1042.8 | 10
C22 FAME | 1357.2 | 5
C22:1 FAME | 1373.3 | 10
C23 Alkane | 1196.8 | 10
C24 FAME | 1522.3 | 10
C25 Alkane | 1350.6 | 5
C26 FAME | 1665.6 | 10
C28 FAME | 1816.6 | 10
C30 FAME | 2000.3 | 10
C31 Alkane | 1786.4 | 10
C32 FAME | 2243.6 | 10
C34 FAME | 2575.6 | 10
C36 FAME | 3034.2 | 10
Ref 1 | 40.1 | 10
Ref 2 | 89.9 | 10
Ref 3 | 139.6 | 10

## Derivatization
comp | class | bound_hydrogen_count | derivative_hydrogen_count | derivative_dD | derivative_dD_uncertainty
C16 | FAME | 31 | 3 | -192 | 3.4
C18 | FAME | 35 | 3
C20 | FAME | 39 | 3
C22 | FAME | 43 | 3
C22:1 | FAME | 41 | 3
C24 | FAME | 47 | 3
C26 | FAME | 51 | 3
C28 | FAME | 55 | 3
C30 | FAME | 59 | 3
C32 | FAME | 63 | 3
C34 | FAME | 67 | 3
C36 | FAME | 71 | 3